JA CC:  CA RDIOVA S CULA R   INTERVENTIONS                                                                                                                                                                                        VOL .   -,   N O .   -,  2 0 2 3

©  2 0 2 3  SOC I ETY   FOR   CAR D I OVAS CU LAR  A NG I O GRAP HY   A ND   I N TE RVE NT I O N S   F O U NDAT I ON ,

H EA RT   RHYTH M  S O CI E TY  AN D  A MERICA N   CO L LEGE   OF   CA RDI O LOGY,   P UBL IS H ED   BY   ELS E V IER IN C .  TH IS   IS  A N   OPE N  A C C E SS  ARTI C L E   U ND ER   TH E   C C   BY - NC - N D   L IC E NSE   ( h ttp:/ /

cre a tivecomm on s . org / licen s es /by - n c - n d/4 . 0/)  WHI C H  P ERMIT S  U S E A N D  DI ST RIBUT I ON  IN AN Y   MED IUM,  P ROVI D E D  T HE  ORIG IN A L W O RK  IS  PROP ERLY  CI TED,  T H E  U S E  IS  N O N - COMMERCIA L A N D N O  MO DI FI CA TI O N S  O R  A D A P T AT I O NS  AR E  M AD E .

|  |  | STANDARDS AND GUIDELINES |  |
| --- | --- | --- | --- |

# SCAI/HRS Expert Consensus Statement on Transcatheter

# Left Atrial Appendage Closure

Jacqueline Saw, MD, FSCAI, Chaira David R. Holmes, MD, FSCAI,

Vice-Chairb

João L. Cavalcante, MDc

James V. Freeman, MD, MPH, MSd

Andrew M. Goldsweig, MD, MS, FSCAIe

Clifford J. Kavinsky, MD, PHD, MSCAIf

Issam D. Moussa, MD, MBA, FSCAIg Thomas M. Munger, MDb

Matthew J. Price, MD, FSCAIh

Mark Reisman, MD, FSCAIi

Matthew William Sherwood, MD, MHS, FSCAIj

Zoltan G. Turi, MD, MSCAIk

Dee Dee Wang, MD, FSCAIl

Brian K. Whisenant, MD, FSCAIm

|  | ABSTRACT |  |  |
| --- | --- | --- | --- |
| Exclusion of the left atrial appendage to reduce thromboembolic risk related to atrial ﬁbrillation was ﬁrst performed surgically in 1949. Over the past 2 decades, the ﬁeld of transcatheter endovascular left atrial appendage closure (LAAC) has rapidly expanded, with a myriad of devices approved or in clinical development. The number of LAAC procedures performed in the United States and worldwide has increased exponentially since the Food and Drug Administration approval of the WATCHMAN (Boston Scientiﬁc) device in 2015. The Society for Cardiovascular Angiography & Interventions (SCAI) has previously published statements in 2015 and 2016 providing societal overview of the technology and institutional and operator requirements for LAAC. Since then, results from several important clinical trials and reg- istries have been published, technical expertise and clinical practice have matured over time, and the device and imaging technologies have evolved. Therefore, SCAI prioritized the development of an updated consensus statement to provide recommendations on contemporary, evidence-based best practices for transcatheter LAAC focusing on endovascular devices. (J Am Coll Cardiol Intv 2023;-:---) © 2023 Society for Cardiovascular Angiography and Interventions Foundation, Heart Rhythm Society and American College of Cardiology, published by Elsevier Inc. This is an open access article under the CC BY-NC-ND license (http://creativecommons.org/licenses/by-nc-nd/4.0/) which permits use and distribution in any medium, provided the original work is properly cited, the use is non-commercial and no modiﬁcations or adaptations are made. |  |  |  |

Endorsement: This statement was endorsed by the American College of Cardiology and the Society of Cardiovascular Computed Tomography.

From the aDivision of Cardiology, Vancouver General Hospital, University of British Columbia, Canada; bDepartment of Cardiovascular Medicine, Mayo  Clinic Health System Rochester, Rochester, Minnesota; cDepartment of Cardiology, Minneapolis Heart Institute, Minneapolis, Minnesota; dSection of Cardiovascular Medicine, Yale University, New Haven, Connecticut; eDivision of Cardiovascular Medicine, The University of Nebraska Medical Center,  Omaha, Nebraska; fDivision of Cardiovascular Medicine, Rush University, Chicago, Illinois; gCarle Heart and Vascular Institute, Carle Illinois College of Medicine, Urbana,  Illinois;  hDivision  of Cardiovascular  Diseases,  Scripps  Green  Hospital,  La  Jolla,  California; iDivision  of Cardiology, Weill  Cornell  Medical Center, New York, New York; jDivision of Cardiology, Inova Heart and Vascular Institute, Fairfax, Virginia; kCenter for Structural and Congenital  Heart Disease, Hackensack University Medical Center, Hackensack, New Jersey; lCenter for Structural Heart Disease, Henry Ford Health System, Detroit,  Michigan; and the mDivision of Cardiology, Intermountain Medical Center, Murray, Utah.

Address for correspondence: Jacqueline Saw. E-mail: jsaw@mail.ubc.ca.

J.L.C. is a Society of Cardiovascular Computed Tomography representative.

J.V.F. and T.M. are Heart Rhythm Society representatives.

M.W.S. is an American College of Cardiology representative.

This article has been copublished in Journal of the Society for Cardiovascular Angiography & Interventions and Heart Rhythm.

ISSN 1936-8798/$36.00                                                                                                                                 https://doi.org/10.1016/j.jcin.2023.01.011

2                 Saw et al                                                                                                                                          J A C C :    C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

| AB BRE V IATIO NS AN D AC RO NYM S DOAC = direct oral anticoagulation DRT = device-related thrombosis LA = left atrium LAA = left atrial appendage LAAC = left atrial appendage closure NCDR = National Cardiovascular Data Registry OAC = oral anticoagulation PVI = pulmonary vein isolation TEE = transesophageal echocardiography TSP = transseptal puncture | ## TABLE OF CONTENTS SUMMARY OF RECOMMENDATIONS . . . . . . . . . . . - METHODOLOGY . . . . . . . . . . . . . . . . . . . - INTRODUCTION . . . . . . . . . . . . . . . . . . . - PATIENT SELECTION FOR LAAC . . . . - Transcatheter LAAC is appropriate for patients with nonvalvular AF with high thromboembolic risk who are not suited for long-term OAC and who have adequate life expectancy (minimum >1 year) and quality of life to beneﬁt from LAAC. There should be patient-provider discussion for shared decision making . . . . . . . . . . . . . . . . . . . . . . . . - |
| --- | --- |

PHYSICIAN AND INSTITUTIONAL

REQUIREMENTS  .  .  .  .  .  .  . . . . . . . . . . . . . . . . . . . . . .  - Physician initial requirements  . . . . . . . . . . . . . . . . . . . .  - Skill maintenance   . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  - Institutional requirements   . . . . . . . . . . . . . . . . . . . . . . .  -

PRE PROCEDURAL IMAGING   . . . . . . . . . . . . . . . . . . . .  -

Baseline preprocedural imaging with TEE or cardiac

computed tomography angiography is

recommended before LAAC . . . . . . . . . . . . . . . . . . . .  -

INTRAPROCEDURAL IMAGING  .  . . . . . . . . . . . . . . . . .  -

Intraprocedural imaging guidance with TEE or ICE and

contrast angiography is strongly recommended  . . .  -

TECHNICAL ASPECTS FOR SAFE AND EFFECTIVE

PROCEDURE PERFORMANCE   .  . . . . . . . . . . . . . . .  -

Procedural venous access, anticoagulation, transseptal

puncture, delivery sheath selection and placement,

LA pressure measurement, and device deployment

should be performed in accordance with the

testing and labeling of each speciﬁc LAAC device . .  -

AVOIDANCE, RECOGNITION, AND TREATMENT OF

INTRAPROCEDURAL COMPLICATIONS  . . . . . . . .  -

Operators need to be familiar with avoidance,

recognition, and management of complications

associated with LAAC   .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  .  -

Predischarge imaging should be performed with 2D

transthoracic echocardiography to rule out

pericardial effusion and device embolization.

Same-day discharge may be appropriate after

several hours of observation, demonstrating no

complications or pericardial effusion after LAAC   . .  -

RECOGNITION AND MANAGEMENT OF

LATE ADVERSE EVENTS . . . . . . . . . . . . . . . . . . . . .  -

Device-related thrombus should be treated with

anticoagulation. Repeat imaging at 45- to 90-day

intervals can be performed to assess for resolution

with eventual cessation of anticoagulation  . . . . . . .  -

Routine closure of iatrogenic atrial septal defects

associated with LAAC should not be performed  . . .  -

The clinical impact and management of PDLs are not

fully understood and all efforts should be made to

minimize such leaks at the time of implantation  . .  -

Patients should be prescribed antithrombotic therapy

with warfarin, direct oral anticoagulants, or DAPT

after LAAC according to the studied regimen and

instructions for use for each speciﬁc device and

tailored to bleeding risks of each patient . . . . . . . . .  -

TEE or cardiac CT is recommended at 45 to 90 days

after LAAC for device surveillance to assess for PDL

and  DRT  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  -

Combined procedures with LAAC (eg, structural

interventions, PVI) are not routinely recommended

pending data from ongoing RCTs . . . . . . . . . . . . . . .  -

FUTURE DIRECTIONS AND LANDSCAPE

OF LA AC    . . . . . .   .  . . . . . .   .  . . . . .   .  . . . . . .   .  . . . . . .   .  -

FUNDING SOURCES   . . . . . . . . . . . . . . . . . . . . . . . . . . .  -

REFERENCES   . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  -

SUPPLEMENTAL MATERIAL   . . . . . . . . . . . . . . . . . . . .  -

## SUMMARY OF RECOMMENDATIONS

1. Transcatheter left atrial appendage closure (LAAC) is appropriate   for   patients   with   nonvalvular   atrial ﬁbrillation  with  high  thromboembolic  risk  who  are not suited for long-term oral anticoagulation and who have  adequate  life  expectancy  (minimum  >1  year) and quality of life to beneﬁt from LAAC. There should be  patient-provider  discussion  for  shared  decision making.

2.1. Physicians  performing   LAAC   should  have   a  prior experience,  including  ≥50  prior  left-sided  ablations or  structural  procedures  and  ≥25  transseptal  punc- tures    (TSPs) .   Interventional    imaging   physicians should have  experience  in guiding  ≥25 TSPs before supporting any LAAC procedures independently.

2.2. For   maintenance   of  skills,   implanting   physicians should perform  ≥25 TSPs and  >12  LAACs  over  each 2-year period.

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

2.3. New  programs  and  implanting  physicians  early  in their  LAAC  experience  should  have  on-site  cardio- vascular surgery backup.

3. Baseline imaging with transesophageal echocardiog- raphy (TEE) or cardiac computed tomography is rec- ommended before LAAC.

4. Intraprocedural imaging guidance with TEE or intra- cardiac echocardiography is recommended.

5. Technical aspects of the procedure, including venous access,  anticoagulation,   transseptal   puncture,   de- livery  sheath   selection   and   placement,   left   atrial pressure   measurement,   and   device   deployment, should be performed in accordance with the labeling of each speciﬁc LAAC device.

6. Operators   need   to   be   familiar   with    avoidance, recognition, and management of procedural compli- cations associated with LAAC.

7. Predischarge  imaging  should  be  performed  with  2- dimensional  transthoracic  echocardiography  to  rule out  pericardial   effusion   and   device   embolization. Same-day discharge may be appropriate after several hours of observation demonstrating no complications or pericardial effusion after LAAC.

8. Device-related   thrombus   should   be   treated   with anticoagulation. Repeat imaging at 45- to 90-day in- tervals can be performed to assess for resolution with eventual cessation of anticoagulation.

9. Routine  closure  of  iatrogenic  atrial  septal  defects associated with LAAC should not be performed.

10. The  clinical  impact  and  management  of peridevice leaks are not fully understood, and all efforts should be  made   to   minimize   such   leaks   at   the   time   of implantation.

11. Patients  should  be  prescribed  antithrombotic  therapy with warfarin,  direct oral anticoagulants, or  dual anti- platelet  therapy  after  LAAC  according  to  the  studied regimen and instructions for use for each speciﬁc device and tailored to the bleeding risks of each patient.

12. TEE   or   cardiac   computed   tomography   is   recom - mended at 45 to 90 days after LAAC for device sur- veillance  to  assess  for  peridevice  leak  and  device- related thrombus.

13. Combined  procedures  with  LAAC  (eg,  structural  in- terventions,   pulmonary    vein    isolation)    are    not routinely  recommended,  as  data  are  pending  from ongoing randomized controlled trials.

## METHODOLOGY

This statement has been developed according to the So- ciety   for   Cardiovascular  Angiography   &   Interventions (SCAI) Publications Committee policies for writing group composition, disclosure and management of relationships

Saw et al                  3

with industry, internal and external review, and organi - zational approval.1

The  writing  group  has  been  organized  to  ensure  di - versity  of perspectives and  demographic  characteristics, multistakeholder representation, and appropriate balance of  relationships  with  industry.  Relevant  author  disclo - sures  are   included   in  Supplemental   Table   S1 .  Before appointment, members of the writing group were asked to disclose ﬁnancial and intellectual relationships from the

12  months  before  their  nomination.  A  majority  of  the writing group disclosed no relevant, signiﬁcant ﬁnancial relationships.   Disclosures   were   periodically   reviewed during  document  development  and  updated  as  needed. SCAI policy requires that writing group members with a current,   relevant   ﬁnancial   interest   are   recused   from participating  in  related  discussions  or voting  on  related recommendations.  The  work  of  the  writing  committee was  supported  exclusively by  SCAI,  a  nonproﬁt medical specialty  society,  without  commercial  support.  Writing group members contributed to this effort on a volunteer basis and did not receive payment from SCAI.

Literature searches were performed by group members designated to lead each section, and initial section drafts were authored primarily by the section leads in collabo - ration  with  other  members  of  the  writing  group.  The recommendations  and  supporting  text  for  each  section were discussed and agreed upon by the full writing group. All recommendations are supported by a short summary of the evidence or speciﬁc rationale. The draft manuscript was peer reviewed in August 2022, and the document was revised   to   address   pertinent   comments.   The   writing group  unanimously  approved  the  ﬁnal  version  of  the document. The SCAI Publications Committee and Execu- tive Committee endorsed the document as ofﬁcial society guidance in December 2022 .

The  SCAI  statements  are  primarily  intended  to  help clinicians  make  decisions  about  treatment  alternatives. Clinicians  also  must  consider  the  clinical  presentation, setting,  and  preferences  of individual  patients  to  make judgments about the optimal approach.

## INTRODUCTION

Atrial  ﬁbrillation  (AF)  is  associated  with  a  4 -  to  5-fold increased  risk  of ischemic  stroke  and  accounts  for  25% of the 700,000 cerebrovascular accidents that occur in the United States annually.2  Historically, the standard of care for stroke prevention in AF has been oral anticoagulation (OAC); however, there are many patient, prescriber, and health  care  resource  issues  that  limit  OAC  use  in  this setting.  Warfarin,  the  traditional  oral  anticoagulant,  is limited by  dietary restrictions, and the need for routine blood  testing  to  maintain  a  narrow  therapeutic  window can lead to patient nonadherence. The newer direct OAC

4                 Saw et al                                                                                                                                          J A C C :    C A R D I O V A S C U L A R    I N T E R V E N T I O N S   V O L .   - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

FIGURE 1 Examples of Transcatheter LAAC Devices

(A) WATCHMAN FLX, (B) Amulet, (C) LAmbre, (D) Wavecrest, and (E) Conformal. LAAC, left atrial appendage closure.

(DOAC)   carries   high   costs,   which   are   not   uniformly absorbed by third party payors. Despite ease of adminis- tration, 30% of patients taking DOAC are nonadherent at 2 years.3   Other   patients   cannot  tolerate  long-term   OAC because of bleeding complications, cognitive impairment, fall  risk,  and  other  factors  (eg,  drug  allergy,  drug  in- teractions, renal dysfunction) . Older patients often have unfavorable bleeding risk proﬁles for OAC, leading phy- sicians not to offer or to discontinue OAC. This treatment gap  has  created  an  unmet  clinical  need  for  an  effective and safe nonpharmacologic therapy for stroke prevention

in patients with nonvalvular AF and has fueled the ﬁeld of LAAC.

In  patients  with  nonvalvular  AF,   >90%  of  thrombi observed  on transesophageal  echocardiography  (TEE)  or on  autopsy  series  come  from  the  left  atrial  appendage (LAA) .4,5   This   ﬁnding   has   stimulated   development   of procedures  to  exclude  the  LAA,  including  endovascular occlusion, surgical suturing, stapling, and amputation, as a  nonpharmacologic  means  to  reduce  stroke  risk.  The LAAOS III trial showed that surgical LAA closure (LAAC) using multiple techniques was associated with reduction

TABLE 1      Ongoing Endovascular LAAC Randomized Controlled Trials and Postprocedural Antithrombotic Strategies

OAC-Eligible Patients

OAC-Contraindicated Patients

| Trial | OPTION Trial: WATCHMAN FLX vs OAC After PV Ablation | CHAMPION-AF Trial: WATCHMAN FLX vs DOAC | CATALYST Trial: Amulet vs DOAC | ASAP-TOO Trial: WATCHMAN vs Control | STROKE-CLOSE Trial: Amulet vs Control | CLOSURE-AF Trial: LAAC vs OAC |
| --- | --- | --- | --- | --- | --- | --- |
| N | 1,600 | 3,000 | 2,650 | 888a | 750 | 1,512 |
| Postprocedural antithrombotic strategies | DOAC or warfarin and aspirin for 3 mo after LAAC | DOAC and aspirin or DAPT for 3 mo after LAAC | DAPT for 3 mo after LAAC | DAPT for 3 mo after LAAC | Aspirin ± clopidogrel for 45 d after LAAC | DAPT after LAAC |
| Control | OAC | DOAC | DOAC | Aspirin or none | OAC, antiplatelet, or none | DOAC or warfarin |

aStudy stopped prematurely (~500 patients enrolled).

DAPT, dual antiplatelet therapy; DOAC, direct oral anticoagulation; LAAC, left atrial appendage closure; OAC, oral anticoagulation; PV, pulmonary vein.

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3                                                                                                                                                                                         Saw et al                  5

- ,  2 0 2 3 : - - -

| TABLE 2 | 2019 AHA/ACC/HRS Focused Update of the 2014 AHA/ACC/HRS Guideline for the Management of Patients With Atrial Fibrillation |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Recommendations |  |  |  | Class of Recommendation | Level of Evidence |
| After surgical occlusion or exclusion of the LAA, it is recommended to continue anticoagulation in at-risk patients with AF for stroke prevention. |  |  |  | I | B |
| ### LAA occlusion may be considered for stroke prevention in patients with AF and contraindications for long-term anticoagulant treatment (eg, those with a previous life-threatening bleed without reversible cause). |  |  |  | IIb | B |
| Surgical occlusion or exclusion of the LAA may be considered for stroke prevention in patients with AF undergoing cardiac surgery. |  |  |  | IIb | B |
| Surgical occlusion or exclusion of the LAA may be considered for stroke prevention in patients undergoing thoracoscopic AF surgery. |  |  |  | IIb | B |
| Adapted from January et al.19 ACC, American College of Cardiology; AF, atrial ﬁbrillation; AHA, American Heart Association; HRS, Heart Rhythm Society; LAA, left atrial appendage. |  |  |  |  |  |
| of ischemic stroke and systemic embolism in patients who continued to receive chronic OAC.6 Whether LAA surgical excision is effective in preventing ischemic stroke in pa- tients not receiving OAC remains unproven. Trans- catheter endocardial devices have been and are actively being evaluated in clinical trials, and the WATCHMAN (Boston Scientiﬁc) and Amulet (Abbott) devices have received Food and Drug Administration (FDA) approval for stroke prevention without long-term OAC. There are several other devices at various stages of development that can be anticipated to add to the procedural arma- mentarium in the future (Figure 1) . Left atrial appendage closure has been compared with OAC for stroke prevention in nonvalvular AF in 3 pro- spective randomized controlled trials (RCTs) showing noninferiority in reduction of stroke and systemic embo- lism: (1) PROTECT AF (707 patients) (NCT001295457) and (2) PREVAIL (407 patients) (NCT011824418) in which the WATCHMAN device was compared with warfarin; and (3) PRAGUE-17 (402 patients) (NCT024269449) in which Amulet and WATCHMAN devices were compared to DOAC, primarily apixaban.10-12 Several important DOAC vs LAAC RCTs for patients eligible for OAC are ongoing, including CHAMPION-AF (NCT0439454613), CATALYST (NCT0422654714), and OPTION (NCT0379529815) (Table 1) . The ASAP-TOO study (NCT0292849716) (stopped prema- turely due to slow enrollment after ~500 patients), STROKE-CLOSE (NCT0283015217), and CLOSURE-AF (NCT0346331718) studies randomize OAC-ineligible pa- tients, but enrollment has been challenged by competi- tion from “off-label” use of LAAC in the OAC-contraindicated population. |  | ## PATIENT SELECTION FOR LAAC 1. TRANSCATHETER LAAC IS APPROPRIATE FOR PATIENTS WITH NONVALVULAR AF WITH HIGH THROMBOEMBOLIC RISK WHO ARE NOT SUITED FOR LONG-TERM OAC AND WHO HAVE ADEQUATE LIFE EXPECTANCY (MINIMUM >1 YEAR) AND QUALITY OF LIFE TO BENEFIT FROM LAAC. THERE SHOULD BE PATIENT-PROVIDER DISCUSSION FOR SHARED DECISION MAKING. Recent guidelines issued by the United States and European professional societies19,20 have both made a IIb recommendation for LAAC in those with AF and contraindication to long-term OAC (Tables 2 and 3) .11,20-23 Because the RCTs were designed to demonstrate noninferiority of LAAC to long-term OAC, they enrolled only patients eligible for long-term OAC.11,21 The mandate by the Centers for Medicare and Medicaid Services (CMS) to perform LAAC for patients for whom long-term OAC is not appropriate had a limited evidence base at the time of the coverage decision. However, the practice in both the United States and worldwide is to treat patients with LAAC who are suboptimal candidates for long-term OAC. CMS also added requirements applied across a range of new technologies, including obligatory registry enrollment and documentation of evidence- based shared decision making with a nonimplanting physician. The National Cardiovascular Data Registry (NCDR) LAA Occlusion (LAAO) Registry reports prior clinical bleeding in almost 70% of patients undergoing LAAC, intracranial bleeding in nearly 12% (vs none in RCTs), and signiﬁcant rates of discharge without OAC (7.9%) in the United States.24,25 In Europe, LAAC without |  |  |  |
| TABLE 3 2020 European Society of Cardiology Guidelines for the Management of Atrial Fibrillation |  |  |  |  |  |
| Recommendations for Occlusion or Exclusion of the LAA |  |  | Class of Recommendation |  | Level of Evidence |
| LAA occlusion may be considered for stroke prevention in patients with AF and contraindications for long-term anticoagulant treatment (eg, intracranial bleeding without a reversible cause) |  |  | IIb |  | B |
| Surgical occlusion or exclusion of the LAA may be considered for stroke prevention in patients with AF undergoing cardiac surgery |  |  | IIb |  | C |

Adapted from Hindricks et al.20

AF, atrial ﬁbrillation; LAA, left atrial appendage.

6

### Saw et al

the use of OAC after procedure is more widespread: in the EWOLUTION Registry (N = 1021), 61.8% were ineligible for OAC.26  In the Amulet Post-Marketing Registry,  82.8% of 1088 patients were contraindicated to OAC.27

Patients suitable for LAAC should ﬁt the deﬁnition of nonvalvular AF and should be at increased risk of stroke with CHA2DS2-VASc of ≥2 (men) or ≥3 (women) as per the ESC  2020  and  American  College  of  Cardiology  (ACC)/ American   Heart   Association   (AHA)   2019   updated   AF guidelines  for  class  I  indication  for  OAC  for  stroke  pre- vention.19,20   Patients   should   have   increased   bleeding risks  (eg,  high  HAS-BLED  score  ≥3)  or  OAC  intolerance, including  prior  bleeding,  fall  risk,  uncontrolled  hyper - tension,  renal  or  liver  failure,  alcohol  use,  concomitant antiplatelet or nonsteroidal agents, high-risk occupations, noncompliance,   labile   international   normalized   ratio, OAC  intolerance/allergy,  and  drug  interactions.  Patients should  have  an  adequate  life  expectancy  (minimum  >1 year)  and  quality  of life  to beneﬁt  from  LAAC.  In  ques- tionable  cases,  frailty  markers  may  help  guide  decision making. There should be patient-provider discussion for shared   decision   making.   LAAC   is   generally   contra- indicated  in  patients  with  known  LAA  thrombus;  im- plantation  in   such   patients   using   embolic   protection devices  is  reported  but  is  inconsistent  with  published indications.28

## PHYSICIAN AND INSTITUTIONAL REQUIREMENTS

2.1.  PHYSICIAN  INITIAL  REQUIREMENTS.  ≥50 prior left- sided  ablations  or  structural  procedures  and  ≥25  trans- septal punctures

2.2.   SKILL   MAINTENANCE.  ≥25   transseptal   punctures and >12 LAACs over 2 years

### 2.3.  INSTITUTIONAL  REQUIREMENTS.  on-site cardiovas-

cular  surgery  (CVS)  program  backup  during  implanter,s early learning curve

The   SCAI/ACC/Heart   Rhythm   Society   (HRS)   LAAC institutional  and  operator  requirements  were  published in  2015,29,30   which  outlined  standards  for  the  operator knowledge  base  and  procedural  skill  sets,  as  well  as institutional requirements required for LAAC. Physicians, educational  requirements  should  include  the  following: knowledge  of AF,  its  clinical  course  and  medical  thera- pies, comprehension of the CHA2DS2-VASc and HAS-BLED scoring systems for risk stratiﬁcation of patients with AF using OAC, indications and management of OAC agents, indications  and  risks  of  various  invasive  surgical  and catheter techniques  for LAAC,  and  awareness  of patient shared decision making.

For  procedural  skills,  physicians  should  have  knowl- edge  of transseptal puncture  (TSP),  LA/interatrial  fossa/ LAA  anatomy   and   associated   intraprocedural   imaging

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

requirements,  and  prior  experience  with  other  percuta - neous endovascular procedures such as, but not limited to, transcatheter mitral valve interventions, atrial septal defect (ASD)/patent foramen ovale closure, or pulmonary vein isolation (PVI) ablation. This should include 50 prior left-sided ablations or structural procedures and at least 25   TSPs.   Industry-sponsored    device-speciﬁc    training (including proctoring) is an expectation as a mechanism for introducing new technologies at the operator,s insti- tution. For maintenance, ≥25 TSPs and >12 LAAOs over 2 years   were   suggested.   Knowledge   and   management experience of complications such as tamponade and de- vice retrieval is required. Educational and procedural skill sets  have  also  been  reviewed  in  the  2015  electrophysi- ology advanced training statement.31

Institutional  requirements  include  availability  of TEE or   intracardiac   echocardiography    (ICE)    operators,    a multidisciplinary  team,   and   an   on-site   CVS  program, although a recent 2020 European Heart Rhythm Associa- tion/European Association of Percutaneous Cardiovascu - lar  Interventions   document   suggested   on-site   CVS   as optional if transfer within an hour is feasible.32  It is rec - ommended that physicians in their  early learning  curve and  starting  a  program  should  have  on-site  CVS.  A  he- modynamic  recording  system  is  necessary  for  pressure assessment.  An  active  institutional  quality  registry  pro- gram  in  the  United  States  is  mandatory.  It  is  recom - mended   that   interventional   imaging   physicians   have experience  in  guiding  ≥25  TSPs  before  supporting  any LAAC procedures independently.33

## PREPROCEDURAL IMAGING

3.  BASELINE  PREPROCEDURAL  IMAGING  WITH  TEE  OR CARDIAC  COMPUTED  TOMOGRAPHY  ANGIOGRAPHY  IS RECOMMENDED  BEFORE  LAAC.  Baseline  preprocedural imaging with TEE or cardiac computed tomography (CT) angiography    (cardiac    CT)    is    recommended    before embarking on the LAAC procedure. The most commonly employed preprocedural LAA imaging is TEE. Some expe- rienced centers have evolved to performing the baseline TEE at the beginning ofthe LAAC procedure. This practice may be  acceptable but  should be  discouraged with  new programs and inexperienced operators. Furthermore, this approach  will  necessitate  periodically  cancelling  a  case with the patient on the table secondary to technical issues such   as   a   lipomatous   or   aneurysmal   septum,   LAA thrombus, or unanticipated technical challenges second- ary to LAA size/morphology. Two-dimensional (2D) images of the LAA are recommended to be obtained at 0。, 45。, 90。, and 135。along the axis of the circumﬂex artery to identify the maximal width of the LAA landing zone.34  Depth di- mensions are  obtained from the  centroid  of the landing zone  as  a  straight  line  to  the  main  body  of  the  LAA

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3                                                                                                                                                                                         Saw et al                  7

- ,  2 0 2 3 : - - -

FIGURE 2 Transesophageal Echocardiogram (TEE) and Computed Tomography Angiography (CTA) Images of the Left Atrial Appendage (LAA) Landing Zone and Deﬁnition of Optimal LAA Length Measurements

(A) TEE images demonstrate the multilobed LAA that emanates from a larger body of the main lobe of the LAA. The landing zone is depicted by an orange line measured at the level of the circumﬂex artery. The length of the main lobe of the LAA is depicted by the blue line measured from the centroid of the LAA landing zone (orange line) to the center of the periapical region of the main body of the LAA. Care is taken to avoid measuring lengths to accessory LAA lobes as this may not accurately mimic the ability of current generation LAA occlusion (LAAO) devices to fully expand within very narrow portions of the distal lobes of the LAA. (B) CTA images demonstrate the multilobed LAA that emanates from a larger body of the main lobe of the LAA. The landing zone is depicted by a red line measured from the centroid of the LAA landing zone (demarcated by solid orange line), to the center of the periapical region of the main body of the LAA. Care is taken to avoid measuring lengths to accessory LAA lobes as this may not accurately mimic the ability of current generation LAAO devices to fully expand within very narrow portions of the distal lobes of the LAA.

(Figure  2) .  To  ensure  that  maximal  dimensions  are  ob- tained,  measurements  should  be  taken  during  maximal ﬁlling of the LAA (left ventricular systole), typically at a mean LA pressure >12 mm Hg.35,36 Three-dimensional (3D) TEE imaging of the LAA evaluated using double-oblique multiplanar reformatting can provide improved accuracy for sizing measurements.

Cardiac CT is increasingly recognized as an excellent 3D primary imaging tool forLAAC procedural planning and is a superior alternative to TEE for spatial resolution. Cardiac CT-gated  scans  should  be  obtained  utilizing  previously described   methodologies   to   ensure   maximal   contrast opaciﬁcation  of the  LAA  and  landing  zone  while  mini - mizing  motion  artifacts.36,37  A  late-pass  delayed  scan  is

8

Saw et al                                                                                                                                          J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

FIGURE 3 Corresponding Images on Fluoroscopy (Showing Intracardiac Echocardiography [ICE] Probe Position), ICE Views, and Corresponding Transesophageal Echocardiography (TEE) Views in the Same Patient

(A) Fluoroscopy showing ICE probe in the retroﬂex position in the LA (also showing the VersaCross wire in the left atrium [LA]), (B) image from the ICE probe in the retroﬂex LA position, (C) image from TEE at 45。view. (D) Fluoroscopy showing ICE probe in the supramitral position in the LA, (E) image from the ICE probe in the supramitral LA position, (F) image from TEE at 0。view. (G) Fluoroscopy showing ICE probe in the left upper pulmonary vein, (H) image from the ICE probe in the left upper pulmonary vein, (I) image from TEE at 90o view.

recommended to allow adequate contrast mixing to avoid overinterpretation  of the  thrombus  in  the  LAA.  Recon- structive segmentation and analysis for device size should be performed at scan increments ideally between 0.5- to 1.5-mm slice thickness with delayed imaging to ensure the absence of LAA thrombus.36  Cardiac CT measurements of the LAA’s landing zone should be performed utilizing 3D multiplanar reformat images of the LAA at the level of the circumﬂex artery in the mid-to-late systolic phase of the cardiac cycle to ensure maximal  LAA contrast  opaciﬁca- tion.38,39  Device sizing should take into account that LAA measurements on cardiac CT tend to be 2- to 3-mm larger than TEE measurements,37,40 given that CT measurements

are performed in 3D and patients are not fasting (and often administered  ﬂuid) before  scanning.  The  addition  of 3D printing and computer simulation to gauge how different device sizes can ﬁt in individual patient anatomies may be helpful, especially with challenging anatomies.

## INTRAPROCEDURAL IMAGING

4.  INTRAPROCEDURAL  IMAGING  GUIDANCE  WITH  TEE OR  ICE  AND  CONTRAST  ANGIOGRAPHY  IS  STRONGLY RECOMMENDED.  It    is    strongly    recommended     that intraprocedural   imaging   guidance   be   utilized   for   all LAAC  procedures.  Intraprocedural  imaging  is  essential

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

### for  ruling  out  LAA thrombus;  monitoring  for  pericardial

effusion;  guidance  of  TSP;  visualization  of  device  im - plantation,    stability,    compression,    and    PDL;     and assessment   of  iatrogenic   ASD.41,42    Fluoroscopy   alone without TEE/ICE is not recommended. TEE guidance by a  dedicated  experienced  interventional  imaging  physi- cian  with  standardized  views  as  described  above  has been  the  traditional  intraprocedural  imaging  modality for  LAAC.  It  is  recommended  that  new  programs  and proceduralists  early in their  LAAC  experience utilize 2D and  3D  TEE  to  become  comfortable  with  imaging  for LAAC.  Intracardiac  echocardiography  is  evolving  as  an additive  imaging  modality  to  guide  LAAC.  The  2D  and emerging  3D  ICE  technology  with  availability   of  pre- procedural   cardiac   CT   case   planning   is   an   emerging acceptable  alternative  to  TEE  for  intraprocedural  guid-

## ance  for  LAAC.  There  is  the  increased  expense  of  the

ICE  probe  that  needs  to  be  considered,  as  well  as  the increased  complexity  of  the  procedure  associated  with ICE  probe  placement  in  the  left  atrium  (LA) .  ICE  im- aging   of  the   LAA   can   be   performed   from   the   right atrium  but  is  usually  suboptimal  for  procedural  guid- ance.  Ideally,  the  ICE  probe  needs  to  be  placed  in  the LA.  This  can  be  performed  through  a  separate  trans- septal  puncture  or  through  the  same  puncture  used  for LAAC  device  delivery  sheath.  ICE  views  (with  probe  in the   LA   in   the   retroﬂex,   supramitral,   and  left   upper pulmonary  vein  positions)  that  simulate  standard  TEE views  may  be  achievable  (Figure  3);  however, in most cases,  TEE  still  provides  superior  imaging  and  operator reproducibility than ICE for the LAA. The RCTs leading to approval of the commercially available LAAC devices both required  intraprocedural  TEE  guidance.  Although  small studies suggest that an ICE-guided approach may be safe and effective, further data in larger cohorts are needed to conﬁrm these early ﬁndings. Therefore, ICE guidance for LAAC, currently, should be reserved for experienced pro- grams/implanters/interventional  imaging  physicians.  In addition, when beginning ICE guidance for LAAC, a period of utilizing combined TEE and ICE is recommended until comfort  with  the   ICE   catheter   and  image   acquisition workﬂow by  the  interventional  imaging  and  implanting team is achieved. Furthermore, when using ICE, baseline preprocedure day imaging with cardiac CT or TEE is rec- ommended to provide detailed information regarding LAA morphology.41,42   Fluoroscopy  alone  without  TEE/ICE  is not recommended.

TECHNICAL ASPECTS FOR SAFE AND

EFFECTIVE PROCEDURE PERFORMANCE

5.    PROCEDURAL   VENOUS   ACCESS,   ANTICOAGULATION, TRANSSEPTAL  PUNCTURE,  DELIVERY  SHEATH  SELECTION AND PLACEMENT, LA PRESSURE MEASUREMENT, AND DEVICE

Saw et al                  9

DEPLOYMENT  SHOULD  BE   PERFORMED   IN   ACCORDANCE WITH     THE  TESTING  AND   LABELING  OF  EACH  SPECIFIC

#### LAAC       DEVICE. LAAC    im plantation   step-by-step.

A  detailed  description  of  procedural  steps  for  LAAC  is beyond  the  scope  of this  article.  Rather,  certain  points requiring clariﬁcation will be highlighted.

Venous access. The  LAAC  procedure  should  generally be performed   through   the   right   femoral  vein   whenever possible. Micropuncture technique and ultrasound guid- ance  can be  helpful  and  should be  strongly  considered. The  left  femoral  vein  may  be  utilized;  however,  trans - septal puncture and subsequent sheath manipulation can be more challenging. After completion of the procedure, hemostasis   can  be   achieved   through   any   number   of techniques    including    manual    compression,    suture- mediated   closure,   or   ﬁgure-of-8   suture   with    equal efﬁcacy     in     accordance     with     operator     comfort. Administration   of   protamine   may   be   considered   if hemostasis is not achieved by standard techniques.

Anticoagulation. Unfractionated heparin (UFH) is used to achieve  therapeutic  anticoagulation  during  the  proced - ure, with a goal activated clotting time (ACT) of at least

250 to 300 seconds. UFH may be administered before or after TSP, at the discretion of the operator. ACT should be monitored    frequently    to    ensure    appropriate    anti- coagulation for the duration of the procedure.

TSP  location. For  the  commercially  available  LAAC  de- vices, the TSP site should be located inferiorly and pos - teriorly. This approach typically orients the access sheath coaxially with  the  LAA.  The  location  of the  LAA  can be conﬁrmed on TEE or preprocedural cardiac CT. For a su- periorly directed LAA, an inferior puncture is desired. In mid-to-inferiorly directed LAA, a less inferior TSP can be performed.   Most   appendages   are   directed   anteriorly;

therefore, a posterior puncture is desired in most cases. On   TEE,   the   90。  view   of  the   LAA   helps   gauge   the direction of the LAA.

Delivery  sheath   selection. Each   commercially   available device comes with manufacturer’s recommendations for speciﬁc  delivery sheaths. For example, the WATCHMAN device can be delivered through 2 different 14F sheaths:

single-curve or double-curve. The double-curve sheath is sufﬁcient to provide coaxial orientation with the LAA in most  cases  where  the  LAA  is  directed  both  superiorly and anteriorly. The single-curve sheath is advantageous when  there   is   an   inferiorly   directed   appendage.   The Amulet   device   can   be   delivered   through   an   inner diameter    12F    (for    devices    ≤25    mm)    or    14F    (for devices  ≥28  mm)  TorqueVue  45  × 45  degree  sheath  or the 14F steerable sheath (outer diameter 19F) .

Advancement of delivery sheath into the LAA. Once TSP is completed, the delivery sheath is introduced into the LA either by anchoring a stiff support wire in the pulmonary vein  or  over  a  “pigtail”  type wire  positioned  in  the  LA.

### 10               Saw et al                                                                                                                                          J A C C :    C A R D I O V A S C U L A R    I N T E R V E N T I O N S   V O L .   - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

| TABLE 4 Procedural and Late Postprocedural Complications of Left Atrial Appendage Occlusion |  |
| --- | --- |
| Periprocedural Complications | Postprocedural Complications |
| Death (<0.2%) Stroke (<0.2%): Ischemic: air or thromboembolism Hemorrhagic Systemic embolism (rare) Pericardial tamponade (~1%) Device embolization (~0.2%) Vascular complications: retroperitoneal bleed, arteriovenous ﬁstula, pseudoaneurysm Other: major bleeding, renal failure, respiratory failure, sepsis, MI, endotracheal/esophageal damage, interfering surrounding structures, device/contrast allergy, pericarditis | #### Late pericardial effusion & tamponade (~1%) Peridevice leak: >5 mm on TEE: 1%-3% >3 mm on TEE: 10%-25% Device-related thrombus (3%-5%) Late device migration/embolization (infrequent) Device erosion (rare) Iatrogenic atrial septal defects (rare to require intervention) |

## MI, myocardial infarction; TEE, transesophageal echocardiography.

Each technique has advantages and disadvantages. A stiff  support wire anchored in the left upper pulmonary vein  provides  an  excellent  support  platform  for  introducing  the device delivery sheath into the LA but runs a greater  risk of wire perforation of the pulmonary vein. The stiff  wire  should be placed in the pulmonary vein through  a  previously  placed  exchange  catheter  positioned  in  the  pulmonary vein using a soft tipped guide wire. A “pigtail” type wire is technically easy and safe to place in the LA  through the transseptal assembly but does not provide as  much  support  for  introduction  of  the  device  delivery  sheath  through   the   TSP.   If  difﬁculty   is   encountered  introducing the  device  delivery  sheath into the LA,  one  can predilate the atrial septum using a 14F Inoue dilator  or  by  performing  atrial  septostomy  using  a  5.0-mm  or  larger   balloon.   Once   in   the   LA,   the   delivery   sheath  should be carefully deaired and ﬂushed with saline. The  device  delivery  sheath  is  advanced  into  the  LAA  over  a  5F to  6F  straight pigtail  catheter positioned  deep in the  LAA.    Commercially    available    “tight    curve”    pigtail  catheters allow for deeper placement in the LAA, a right  anterior   oblique/caudal   projection   can   be   helpful   in  proﬁling the LAA for sheath advancement.

LA pressure measurement. Measuring LA pressure can be helpful in conﬁrming position in the LA during TSP and ensuring that the LA pressure is high enough to properly measure LAA dimension during the procedure to ensure proper device size selection and to avoid choosing a de- vice  too   small   and   risking   embolization.  A   mean   LA pressure of ≥12 mm Hg is recommended for proper LAA measurements  and  device  size  selection.43   Intravenous ﬂuid should be administered when patients arrive for the procedure unless contraindicated, and hydration should be administered until the goal LA pressure is achieved.

Deployment of LAAC device. Operators should refer to the manufacturer’s  instructions  for  use  for  proper   device deployment  for  each  commercially  available  device  and should  adhere  to  established  criteria  to  ensure  proper

positioning before device release. Care should be taken to ensure   proper   device   sizing,   anchoring,   positioning within the LAA, and absence of signiﬁcant PDLs.

AVOIDANCE, RECOGNITION, AND TREATMENT OF INTRAPROCEDURAL COMPLICATIONS

6.      OPERATORS     NEED     TO     BE     FAMILIAR     WITH AVOIDANCE,   RECOGNITION,   AND   MANAGEMENT   OF COMPLICATIONS        ASSOCIATED         WITH          LAAC.

WATCHMAN, the ﬁrst LAAC device approved for general use,  underwent  a  lengthy,  multistepped  approval  con- sisting of 2 randomized clinical trials, 2 continued access registries,  and  an  unprecedented  3  FDA  panel  hearings spanning  2  decades  before  approval  for  commercializa - tion was obtained. This was largely due to safety concerns surrounding  the  LAAC  procedure.  In  the  PROTECT  AF RCT, there was an 8.7% complication rate driven largely by  a  4.8%  serious  pericardial  effusion,  as  well  as  0.6% device embolization, 0.6% air embolism, 0.9% procedure- related stroke, and 0.6% access-site bleeding.21  Following procedure  reﬁnements  and  increasing  operator  experi - ence, the complication rates have  dropped signiﬁcantly. Contemporary  clinical  data  suggest  that  serious  compli- cation   rates   and   pericardial   effusions/tamponade   are well <2% . Despite this, it is essential that all implanters be skilled  in  the  recognition  and  treatment  of  procedural complications and be attentive to the details necessary to prevent such events (Table 4) .

Pericardial  effusion. Imaging  of the  pericardial  space should be performed at the beginning and  end  of every LAAC  procedure  to  assess  for  pericardial  effusion.  Sys - temic blood pressure should be monitored throughout the procedure with a manual cuff or arterial line to aid prompt diagnosis of tamponade. Pericardial  effusion is typically caused by perforation of the LAA and can result in tam- ponade,  requiring  urgent  pericardiocentesis  and,  rarely, urgent surgery. Pericardial effusion may also be related to

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

### the TSP or damage to other cardiac structures. Increasing

experience with LAAC, preprocedural imaging and plan- ning, the use of intraprocedural echocardiographic guid- ance,  and  utilization  of less  traumatic  guide  wires  and pigtail  catheters  have  resulted  in  a  marked  decrease  in pericardial effusion. In particular, the mandated use of a pigtail catheter for entry into the LAA allowed advance- ment of the large sheath more safely, preventing damage to the thin LAA wall. LAAC has seen a stepwise decrease in serious pericardial effusion from 5.0% to the current level of <1.5% . In the NCDR LAAO Registry of 38,158 patients treated with the commercial WATCHMAN device from2016 to  2018,  pericardial  effusion  requiring  intervention  was documented in 528 (1.39%) patients, of which 437 (83%) were  treated  with  percutaneous  drainage,  primarily  for tamponade.  There  were  91  (17%)  patients  who  required open cardiac surgery.24

Operators  and  other  members  of the  procedure  team should   be   adept    in    recognition    and    prompt    peri- cardiocentesis for effusions associated with any signs of hemodynamic    compromise.    Pericardiocentesis    trays should  be  in  the  room,  and  surgical  teams  should  be immediately   available,    should    surgical    intervention become necessary. If bleeding is still brisk into the peri- cardium  after   pericardiocentesis   pigtail   insertion   and draining, protamine may be considered. Autotransfusion and  surgical  intervention  are  the  immediate  next  man - agement steps. Occasionally, pericarditis may result from pericardial  drainage,  which  is  conventionally  managed with    a    short    course    of    colchicine    and/or    anti- inﬂammatory agents.

Periprocedural stroke. The second major complication identiﬁed  in  the  PROTECT  AF  trial  was  periprocedural stroke,  which  has  decreased  substantially  with  experi- ence. In the NCDR LAAO Registry, an ischemic stroke was documented  in  0.12%,  whereas  hemorrhagic  stroke  was identiﬁed  in  0.01% .24  Several  potential  causes  of stroke have   been   identiﬁed:   (1)   preprocedural   presence   of thrombus  or  sludge;  (2)  air  embolization,  typically  the result of inadequate purging and ﬂushing of equipment; and  (3)  development  of thrombus  on  equipment  during the   procedure.   Rigorous   screening   to   rule   out   LAA thrombus  before  procedure  and  meticulous  ﬂushing  to avoid air embolization are essential. The incidence of air embolism is extremely rare in contemporary practice with careful  preparation  of  equipment.  In  the  presence  of spontaneous  echo   contrast,   operators   should   proceed with  caution  after  ruling  out the presence  of thrombus. Administering  intravenous  UFH   during  the  procedure aiming for ACT 250 to 300 seconds is necessary to reduce the  risk  of thrombus  formation  on  catheters,  wires,  de- vices,    and    so    forth.    Inadequate    procedural    anti- coagulation, prolonged procedure duration, or excessive recaptures of the device (resulting in tissue trauma) may

Saw et al                 11

promote  thrombus  formation.  Once  thrombus  is  identi - ﬁed,  additional  heparin  should  be  administered,  and  a prompt decision should be made whether to continue the procedure after documenting successful resolution of the thrombus  or  to  abort  and  remove  the  equipment  with adherent thrombus.

#### Device embolization . This complication is uncommon.

In  the  NCDR  LAAO  Registry,  30  (0.07%)  embolizations were  reported.24  A  systematic  review  of cases  reported with both the WATCHMAN device and Amplatzer Cardiac Plug identiﬁed 31 cases.44 Two-thirds occurred during the implant,  whereas  the  others  occurred  later,  sometimes detected at follow-up TEE or  cardiac  CT angiography as incidental  ﬁndings  in  patients who were  asymptomatic. Embolized devices may be lodged in the aorta, LA, or left ventricle.  In  21  of the  cases  for  which  information  was available, percutaneous removal was attempted in 17 and a   surgical   approach   attempted   in   10  to   retrieve  the embolized  devices.  The  retrieval  approach  was  in  large part dependent on the location of the embolized device; surgery  was  typically  the  ﬁrst  choice  if the  device  had embolized to the left ventricle and was entangled in the mitral   apparatus   and   associated   with   hemodynamic deterioration.  LAAC  devices  that  had  embolized  to  the descending aorta or LA were often successfully retrieved percutaneously. Although the case numbers were small, potential  mechanisms  for  embolization  included  device undersizing or oversizing. Implanters should be skilled in the use of large sheaths and snares for retrieving devices from the LA or aorta.

Other  complications . The  NCDR  LAAO  Registry  docu- mented    other    important    procedural    complications, namely in-hospital major bleeding that occurred in 1.25% and major vascular  complications in  0.15% .24  In this re- gard, 70% of the patients had a history of prior bleeding, including   42%   with    prior    gastrointestinal   bleeding.

Venous  bleeding  with  the  large  sheath  size  may  occur. Access-site  bleeding  can  be  avoided  by  practicing  good technique  for  gaining  access  and  achieving  hemostasis. Using  ultrasound-guided  venous  puncture  with  micro- puncture   needle/catheter   may   enhance   safety.   Using suture-mediated   preclosure   or   ﬁgure-of-8   suture   or combining the 2 ensures complete hemostasis and mini - mizes the possibility of late access-site bleeding. Avoid - ance   of  excessive   anticoagulation   and   adherence   to standard    hemostasis    techniques    can    help    prevent bleeding.

7.   PREDISCHARGE   IMAGING  SHOULD  BE   PERFORMED WITH   2D   TRANSTHORACIC   ECHOCARDIOGRAPHY   TO RULE     OUT      PERICARDIAL      EFFUSION     AND      DEVICE EMBOLIZATION.      SAME-DAY       DISCHARGE       MAY      BE APPROPRIATE  AFTER  SEVERAL  HOURS  OF  OBSERVATION, DEMONSTRATING   NO   COMPLICATIONS   OR   PERICARDIAL

### 12

Saw et al

FIGURE 4 Device-Related Thrombus (TEE)

TEE, transesophageal echocardiography.

EFFUSION  AFTER  LAAC.  Imaging   with   2D   transthoracic echocardiogram     is     recommended     before     hospital discharge  to  evaluate  for  new  pericardial  effusion  and rule out device embolization. After LAAC, patients were observed overnight in-hospital in the US pivotal trials of the  WATCHMAN  device,  and  this  is  currently  common practice  after  LAAC.  “Inpatient  status”  is  mandated  by CMS for reimbursement.23  Several small case series have evaluated  the   feasibility   of  same-day   discharge  after endocardial  LAAC  and  found  that  in  selected  patients without  procedural  complications,  same-day  discharge can  be  done  without  an  associated  increase  in  adverse events.45,46  More  data  should be  collected before  same - day discharge can be routinely advocated, although that practice  is  increasing  in  frequency  and  may  be  appro- priate after  several hours  of observation,  demonstrating no  complications  or  pericardial  effusion.

RECOGNITION AND MANAGEMENT OF LATE ADVERSE EVENTS

8.  DEVICE-RELATED  THROMBUS  SHOULD  BE  TREATED WITH ANTICOAGULATION.  REPEAT  IMAGING AT 45 - TO 90-DAY  INTERVALS  CAN  BE  PERFORMED  TO  ASSESS FOR   RESOLUTION   WITH   EVENTUAL   CESSATION   OF ANTICOAGULATION.  Device- related        thrombosis . Device-related thrombosis  (DRT)  (Figure 4) has been  re- ported after 3% to 5% of LAAC procedures in large clinical trial and registry experiences.10,11,24  DRT has been shown to be more likely in patients with prior stroke, large LAA dimensions,   those    with    permanent   AF,    and    those with   previous   LA   thrombus.11,47-49     The   relationship between   DRT   and   long-term   adverse   cardiovascular events is unclear, but registry-based analyses have found an    association    between    DRT    and    stroke/transient

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

ischemic  attack  (TIA)  during  short-term  follow-up.10,47

## However,   most   strokes   after   LAAC   occur  in   patients

without DRT; thus, the clinical importance of DRT is not well  understood.11,46   If DRT  is  diagnosed  on  post-LAAC imaging,  anticoagulation  should  be  continued.  For  pa - tients on dual antiplatelet therapy (DAPT), OAC should be started  and  DAPT  should be  discontinued.50  Repeat im - aging  at  45-  to  90-day  intervals  can  be  performed  to assess  for  resolution  with  eventual  cessation  of  anti - coagulation.  Both  cardiac  CT  and  TEE  have  comparable accuracy for DRT detection51,52  and should be chosen ac - cording  to  site  experience  and  patient’s  preference  and clinical proﬁle.

### 9.  ROUTINE  CLOSURE  OF  IATROGENIC ATRIAL  SEPTAL DEFECTS  ASSOCIATED  WITH   LAAC  SHOULD   NOT   BE

#### PERFORMED.  Residual   iatro ge nic  ASD . Endovascular

LAAC requires traversal  of the interatrial septum with a large  sheath;   subsequent   removal   of  the   sheath  may result  in   left-to-right   or   even  bidirectional   shunting, which  typically  resolves  by  6  months  in  most  cases.53 There  has  been  very  little  study  of  the  clinical  signiﬁ - cance of iatrogenic ASD from the LAAC procedure. Small randomized trials have not shown any clinical beneﬁt in closure of the iatrogenic ASD after transcatheter edge-to- edge  repair  of  the  mitral  valve  or  transcatheter  mitral valve    replacement,    which    require    >20F    sheaths/ devices.54   At  present,  there  is  no  clear  indication  for closure of the iatrogenic ASD associated with LAAC unless there   is    signiﬁcant    right    ventricular    dilation    and/ or dysfunction.

10. THE CLINICAL  IMPACT AND  MANAGEMENT OF  PDLs ARE   NOT   FULLY   UNDERSTOOD   AND   ALL    EFFORTS SHOULD  BE  MADE  TO  MINIMIZE  SUCH  LEAKS  AT  THE TIME  OF   IMPLANTATION.  Peridevice  leak. Peridevice

leaks (PDLs) are common and are reported in 11% to 57% of  transcatheter   LAAC   procedures   depending   on   the implanted   device   and   imaging   modality   used.55-58    A relatively  small  secondary  analysis  of the  PROTECT  AF trial showed no difference in adverse event rates between leak sizes of <1, 1 to 3, and  >3 mm but did not evaluate larger  leaks  with  the  WATCHMAN  device.  Subsequent studies  have  corroborated  these  results,  although  these were  small and may have been underpowered to  detect differences  in  outcomes.58,59   An  analysis  of  51,333  pa- tients in the NCDR LAAO Registry found that a small leak (<5  mm)  was  associated  with  a  slightly  higher  rate  of stroke, TIA, and systemic embolism over 1 year (adjusted hazard ratio, 1.152; 95% CI, 1.025-1.294) .60

There are limited data and no consensus on the man- agement of PDL. PDLs of >5 mm are believed to be sig - niﬁcant,  and  current  practice  is  generally  to  continue OAC.    Case    reports     and     small     case     series     have

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

### demonstrated the feasibility of closing leaks using inter-

ventional approaches including placement of occlusion or plug devices or coiling, but data are lacking on the long- term   impact   of   these   approaches   including   adverse event rates and the safety of discontinuing OAC.61-64

#### Late pericardia l effusion. Late pericardial effusion has

been reported infrequently via case reports and in 1% of patients in the PRAGUE-17 trial.12,65  These may be related to the hooks of the devices penetrating through the LAA wall. Management is similar to periprocedural pericardial effusion.

#### Rare   complications    (device    erosion,    infection,

nickel  allergy) . Left  atrial  appendage  closure  may  be associated with device erosion in rare cases. Erosion can result in damage to the pulmonary artery or late pericar- dial   effusion.15,66    Infection   is   exceedingly   rare   with LAAC67; nevertheless, the instructions for use for both the WATCHMAN and Amulet devices recommend appropriate endocarditis  prophylaxis  for  6  months  after  device  im- plantation.  Treatment  for  device  infection  may  require antibiotics and surgical removal of the device.13  Devices available for LAAC in the United States are predominantly made of nitinol, an alloy of roughly equal parts nickel and titanium. Allergic reactions to nickel have been reported in patients undergoing patent foramen ovale/ASD closure and  usually  present  as  dermatitis,  although  reports  of uncomplicated implantation in patients with documented nickel allergies have also been published.14,68  In general, manufacturer  indications   for  use   recommend   against implantation  of nitinol  devices in patients with  a  docu - mented signiﬁcant nickel allergy.

11.  PATIENTS  SHOULD  BE  PRESCRIBED  ANTITHROMBOTIC THERAPY WITH WARFARIN, DIRECT ORAL ANTICOAGULANTS, OR DAPT AFTER LAAC ACCORDING TO THE STUDIED REGIMEN AND  INSTRUCTIONS  FOR  USE  FOR  EACH  SPECIFIC  DEVICE AND   TAILORED   TO   BLEEDING   RISKS   OF   EACH   PATIENT.

The FDA approved the WATCHMAN device based on the PROTECT AF and PREVAIL trials, in which warfarin and aspirin were administered for 45 days, and if no DRT or PDL of >5 mm was observed on TEE, patients transitioned to DAPT with aspirin and clopidogrel until 6 months after implant and to aspirin alone afterward.12,24,69,70 However, following FDA approval, the NCDR LAAO Registry showed utilization of varied antithrombotic regimens after LAAC in the United States. In this Registry, discharge after LAAC with WATCHMAN device on DOAC and aspirin was asso- ciated with similar outcomes compared with warfarin and aspirin at 45 days and 6 months. Furthermore, discharge on  either  warfarin  or  DOAC  alone  without  aspirin  was associated with a lower risk of adverse events, largely due to lower rates of bleeding.24  In the PINNACLE FLX study, patients were discharged on DOAC and aspirin for 45 days and  then   transitioned   to   DAPT;  there   are   similar   or

Saw et al                13

favorable  outcomes  in  other  studies  compared  with  the older  pivotal  trials.71    The  WATCHMAN  FLX  FDA  label initially allowed for either warfarin or DOAC in addition to aspirin  for  the  initial  45  days  after  device  implantation before  transitioning  to  DAPT.  The  revised  FDA  labeling now allows the use of DAPT for 45 days after WATCHMAN FLX implantation, following a recent propensity-matched study comparing DAPT versus either a DOAC or warfarin plus  aspirin  after  WATCHMAN  FLX  in  the  NCDR  LAAO Registry,  which   showed   no   difference   in  the   rate   of composite  death,  stroke,  or major bleeding, and  device- related thrombus.72

The  use  of postprocedural  DAPT  (without  OAC)  after LAAC is common practice in Europe and other parts of the world, as most patients treated with LAAC have  contra - indications to OAC.68  In the Amulet IDE study, 75.7% of patients were discharged on DAPT after Amulet implan- tation,  and  the  rate  of DRT  was  not  different  from  the WATCHMAN   arm,   in   which   95.8%   were   on   OAC   at discharge.73    Other   small   trials   and   single/multicenter prospective       observational       registries       using       the WATCHMAN and Amplatzer/Amulet devices have shown similar   results.26,49,55,74    The   EWOLUTION   registry   of WATCHMAN  procedures  performed  in  Europe  reported that  60% of patients were  discharged  on DAPT, with an adverse events rate comparable with that published in the pivotal WATCHMAN trials.26,49  The NCDR LAAO Registry showed   that   discharge   on   DAPT   after   WATCHMAN implant  was  relatively  uncommon  (5%  of implants)  but was  associated  with   a   similar  risk   of  adverse   events compared  with  warfarin  and  aspirin  at  45  days  and  6

months.24  A large  meta-analysis  of 83  studies  including

12,326  LAAC patients  (7900  on  antiplatelet,  and 4151  on OAC)  showed no  difference in  stroke,  DRT, bleeding,  or overall   mortality   with   antiplatelet   versus   OAC   after LAAC.75     Currently    enrolling   randomized    trials    will continue   to   evaluate   optimal   postprocedural   antith- rombotic strategies (Table 1) .

In  summary,  the  studies  till  date  are  insufﬁcient  to deﬁne     an     optimal     postprocedural     antithrombotic regimen  following  LAAC.  Discharge  on  anticoagulation with  anticoagulant  plus  aspirin  is  well  established  for WATCHMAN  based  on  the  PROTECT  AF,  PREVAIL,  and PINNACLE FLX studies. Discharge on DAPT appears well established  for  the  Amulet  device  based  on  the  Amulet IDE  study  and  real-world  registries.  There  is  also  an increasing  body  of  registry  evidence  supporting  DAPT with   WATCHMAN.   At   this   juncture,   the   choice   and duration  of  antithrombotic  regimen  after  LAAC  should be  made   according   to   the   studied  regimen   for   each speciﬁc  device,  potentially  individualized  based  on  the patient,s    bleeding/thrombotic    risk,    and    potentially modiﬁed  if  PDL  or  DRT  is  seen  on  post-LAAC  imaging surveillance.

### 14              Saw et al                                                                                                                                          J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S   V O L .   - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

| TABLE 5 Imaging Surveillance Modality and Optimal Imaging at Different Postdevice Implantation Time Points |  |  |  |  |
| --- | --- | --- | --- | --- |
| Imaging Timing | Immediate Postdevice Implant | Prehospital Discharge | 45-d Follow-Up | 1-y Follow-Up (optional) |
| Transthoracic echocardiogram | - | +++ | - | - |
| Transesophageal echocardiogram | +++ | - | ++ | ++ |
| CCTA | - | - | +++ | +++ |
| Complication surveillance | Pericardial effusion | Device embolization | Peridevice leak | Device-related thrombus |
| Transthoracic echocardiogram | +++ | + | - | - |
| Transesophageal echocardiogram | +++ | +++ | ++ | +++ |
| CCTA | +++ | +++ | +++ | +++ |

+++, strongly recommended; ++, less strongly recommended; +, recommended; -, not required. CCTA, cardiac computed tomography angiography.

### 12. TEE OR CARDIAC CT IS RECOMMENDED AT 45 TO 90 DAYS   AFTER   LAAC   FOR   DEVICE   SURVEILLANCE   TO

ASSESS FOR PDL AND DRT.  There is variability in clinical practice on postprocedural imaging modality and timing. Typically,  imaging  surveillance  after  LAAC  device  im- plantation is recommended at 3 time points:

1. Immediately  after  LAAC implant to  evaluate for peri- cardial effusion, device stability, and PDL (with TEE or ICE)

2. Prehospital  discharge  to  evaluate  for  new  pericardial effusion   and   device   embolization   (with   2D   trans- thoracic echocardiography)

3. At 45  days after LAAO  device implant to evaluate for sealing  of the  LAAC  device  or  continued  presence  of PDL.

Common imaging modalities utilized at these speciﬁc time points are depicted in Table 5 .

Of note, some centers delay the 45-day imaging to 90 days, and some centers routinely perform a second TEE/ cardiac  CT  at  1  year  as  per  the  PROTECT  AF  trial  pro- tocol.  The  instructions  for  use  for  WATCHMAN  recom - mends   both   45-day   and   1-year   device   surveillance. During    the    COVID-19    pandemic,    some    physicians migrated    away    from    TEE,    which    generates    aero- solization, to less invasive surveillance imaging with the use  of cardiac  CT.76   There  appears  to  be  greater  sensi- tivity  and  reproducibility  of  cardiac  CT  for  the  evalua- tion  of LAAC  PDL  and  similar  rate  of detection  of DRT compared with TEE.51,52  These reasons together with the noninvasiveness led to many preferring cardiac CT over TEE for  device surveillance  (Table 5) . However, imaging surveillance  after  LAAC  is  an  evolving  topic  and  more comparative  and   prospective   data   are   required.51    We recommend   device   surveillance   with   TEE/CT   angiog- raphy  at  45  to  90  days  after  LAAC,  and  a  second  im- aging at 1 year should be considered if a PDL is there at prior  imaging  or  if there  are  concerns  for  DRT  risk.

13.     COMBINED     PROCEDURES     WITH      LAAC      (EG, STRUCTURAL      INTERVENTIONS,      PVI)      ARE      NOT ROUTINELY    RECOMMENDED    PENDING    DATA    FROM ONGOING    RCTs.  Left    atrial    appendage    closure    is increasingly performed simultaneously with PVI ablation and other structural heart procedures. In a report of 139 patients  from  10  European  centers  who  underwent  AF ablation  (chieﬂy  PVI)  followed by  WATCHMAN  LAAC  in the same setting,77  a high degree of LAAC success (100%) with  excellent  periprocedural  safety  and  an  early  (28 days) 98% freedom from PDL of >5 mm was shown. There were, however, leaks of <5 mm, potentially the result of small  device  migration  with  resolution  of tissue  edema from  PVI.  Good  results  were  conﬁrmed  at  the  2-year follow-up, with 92% off OAC and a relative 84% cerebro - vascular accident/TIA reduction compared with predicted in patients with similar risk scores not treated with OAC.78  The OPTION trial randomized ~1600 patients 1:1 to either OAC  or  WATCHMAN  FLX  following  AF  ablation,  with approximately one-third of the WATCHMAN devices be- ing   placed   concomitantly   immediately   after   ablation. Results   are   anticipated   in   2023 .   The   aMAZE   study (NCT0251379779) prospectively randomized patients 2:1 to percutaneous  suture  closure  of  the  LAA  (LARIAT,  Sen - treHEART) as an adjunct to PVI ablation in patients with symptomatic  or  longstanding  persistent  AF.  This  study showed  that  although  performing  LAA  ligation  together with PVI was safe, there was no reduction in recurrent AF compared with PVI alone.80

The  WATCH-TAVR  trial  (NCT0317353481)  randomized 350 patients with aortic stenosis and AF to transcatheter aortic  valve  replacement  and  medical  therapy  vs  trans - catheter   aortic   valve   replacement   with   concomitant WATCHMAN LAAC with a primary end point of all-cause mortality,  stroke,  or  major  bleeding.  Results  are  antici- pated  in  2023 .  Until  these  RCT  results  are  available, combined procedures with LAAC should not be routinely performed.

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

## FUTURE DIRECTIONS AND LANDSCAPE OF LAAC

The  ﬁeld  of  transcatheter  LAAC  has  evolved  tremen - dously in its ﬁrst 2 decades, with marked improvements in  safety,  efﬁcacy,  and  device  iterations.  Ongoing  trials and  evolving   practices   indicate   signiﬁcant   future   ad - vancements to come. The FDA approval of WATCHMAN FLX  and  Amulet  devices,  in-progress  RCTs  evaluating new  devices,  and  an  array  of other  technologies  in  the development  and   approval   queue  will   mean   a  larger toolbox for LAAC. Several RCTs comparing LAAC to con - trols   in   OAC-ineligible   patients   and   RCTs   comparing DOAC to LAAC in OAC-eligible patients are ongoing and are  anticipated  to  broaden  the  clinical  indications  and strengthen   societal   recommendations   for   this   device therapy.  Advancements  in  periprocedural  imaging  have also improved the safety and efﬁciency of the procedure, and  novel  imaging  technologies  are  continuing  to  be explored. This document provides the current evidence- based best practices of endovascular LAAC by consensus of the  established  expert  panel  with  multisocietal  sup- port.  Given  the  rapid  developments  in  this  ﬁeld,  this

Saw et al                15

consensus   document    acknowledges    important    areas where   further   evidence   is   needed   to   be   developed including but not restricted to the type and  duration  of post-LAAC   antithrombotic  therapy,   standardization   of post-LAAC imaging surveillance and time points; duration of  OAC   after   DRT   detection   and   follow-up   imaging; management  of  PDL;  and  head-to-head  comparison  of LAAC vs DOAC in OAC-eligible patients.

## FUNDING  SOURCES This  research  did  not  receive  any

speciﬁc grant from funding agencies in the public, com - mercial, or not-for-proﬁt sectors.

## DECLARATION OF COMPETING INTEREST

João   Cavalcante  received   consulting  fees   from   Boston   Scientiﬁc   and Abbott.  Matthew  Price  serves  on  the  advisory  board  and  has  received consulting, honoraria, and Speakers Bureau fees from Abbott. Brian Whi - senant received consulting fees from Boston Scientiﬁc and Abbott. Dee Dee Wang  is  a  principal  investigator  for  investigator-initiated  research  for Boston Scientiﬁc and has received consulting fees from Abbott. Jacqueline Saw, David Holmes, Andrew Goldsweig, James Freeman, Thomas Munger, Issam  Moussa,  Clifford  Kavinsky,  Matthew  Sherwood,  Zoltan  Turi,  and Mark Reisman reported no ﬁnancial interests.

## RE FE RE NCES

1. Szerlip M, Feldman DN, Aronow HD, et al. SCAI

publications committee manual of standard operating

procedures. Catheter  Cardiovasc  Interv. 2020;96(1):

145-155.

2. Benjamin  EJ, Muntner P, Alonso A, et al.  Heart

disease and stroke statistics-2019 update: a report

from  the  American  Heart  Association.  Circulation.

2019;139(10):e56-e528.

3. Martinez    C,    Katholing    A,    Wallenhorst    C,

Freedman SB. Therapy persistence in newly diagnosed

non-valvular atrial ﬁbrillation treated with warfarin or

NOAC. A cohort study. Thromb Haemost. 2016;115(1):

31-39.

4. Blackshear JL, Odell JA. Appendage obliteration to

reduce stroke in cardiac surgical patients with atrial

ﬁbrillation. Ann Thorac Surg. 1996;61(2):755-759.

5. Johnson WD,  Ganjoo AK, Stone CD, Srivyas  RC,

Howard M. The left atrial appendage: our most lethal

human attachment! Surgical implications. Eur J Car-

diothorac Surg. 2000;17(6):718-722.

6. Whitlock RP, Belley-Cote EP, Paparella D, et al. Left

atrial appendage occlusion during cardiac surgery to

prevent stroke. N  Engl  J  Med. 2021;384(22):2081-

2091.

7. WATCHMAN  Left  Atrial  Appendage  System  for Embolic PROTECTion in Patients With Atrial Fibrilla- tion.   ClinicalTrials.gov   identiﬁer:   NCT00129545. Accessed July 20, 2021. https://clinicaltrials.gov/ct2/ show/NCT00129545

8. Evaluation   of   the   WATCHMAN   Left   Atrial Appendage  (LAA)  Closure  Device  in  Patients  With Atrial Fibrillation Versus Long Term Warfarin Therapy (PREVAIL). ClinicalTrials.gov identiﬁer: NCT01182441. Accessed July 20, 2021. https://clinicaltrials.gov/ct2/ show/NCT01182441

9. Left  Atrial  Appendage  Closure  vs.  Novel  Anti- coagulation Agents in Atrial Fibrillation (PRAGUE-17). ClinicalTrials.gov identiﬁer: NCT02426944. Accessed July  20,  2021.  https://clinicaltrials.gov/ct2/show/ NCT02426944

10. Holmes DR, Reddy VY, Turi ZG, et al. Percutaneous closure of the left atrial appendage versus warfarin therapy for prevention of stroke in patients with atrial ﬁbrillation: a randomised non-inferiority trial. Lancet. 2009;374(9689):534-542.

11. Holmes DR, Kar S, Price  MJ, et al.  Prospective randomized evaluation of the Watchman Left Atrial Appendage  Closure  device  in  patients  with  atrial ﬁbrillation  versus  long-term  warfarin  therapy:  the

PREVAIL trial. J Am Coll Cardiol. 2014;64(1):1-12.

12. Osmancik P, Herman D, Neuzil P, et al. Left atrial appendage closure versus direct oral anticoagulants in high-risk patients with atrial ﬁbrillation. J  Am  Coll Cardiol. 2020;75(25):3122-3135.

13. WATCHMAN FLX versus NOAC for embolic pro- tection  in the  management of  patients with  non- valvular atrial ﬁbrillation. ClinicalTrials.gov identiﬁer: NCT04394546.  Accessed  July  20,  2021.  https:// clinicaltrials.gov/ct2/show/NCT04394546

14. Amplatzer Amulet LAAO vs. NOAC (CATALYST). ClinicalTrials.gov identiﬁer: NCT04226547. Accessed July  20,  2021.  https://clinicaltrials.gov/ct2/show/ NCT04226547

15. Comparison  of  anticoagulation  with  left  atrial appendage closure after AF ablation (OPTION). Clin- icalTrials.gov identiﬁer: NCT03795298. Accessed July

20, 2021. https://clinicaltrials.gov/ct2/show/NCT037 95298

16. Assessment of the WATCHMAN Device in Patients Unsuitable  for  Oral  Anticoagulation   (ASAP-TOO).

ClinicalTrials.gov identiﬁer: NCT02928497. Accessed

July   21,   2021.   https://clinicaltrials.gov/ct2/show/ NCT02928497

17. Prevention of Stroke  by  Left Atrial Appendage Closure in Atrial Fibrillation Patients After Intracere- bral     Hemorrhage.    ClinicalTrials.gov     identiﬁer: NCT02830152.     https://clinicaltrials.gov/ct2/show/ NCT02830152

18. Left Atrial Appendage CLOSURE in Patients With Atrial  Fibrillation  Compared  to  Medical  Therapy (CLOSURE-AF).        ClinicalTrials.gov        identiﬁer: NCT03463317.  Accessed  July  20,  2021.  https:// clinicaltrials.gov/ct2/show/NCT03463317

19. January CT, Wann LS, Calkins H, et al. 2019 AHA/ ACC/HRS focused update of the 2014 AHA/ACC/HRS guideline for the management of patients with atrial ﬁbrillation: a report of the American College of Car- diology/American  Heart  Association  task  force  on clinical practice guidelines and the Heart Rhythm So- ciety in collaboration with the Society of Thoracic Surgeons.  Circulation.  2019;140(2):e125-e151.  Pub- lished correction appears in Circulation. 2019;140(6): e285.

20. Hindricks G, Potpara T, Dagres N, et al. 2020 ESC guidelines for the diagnosis and management of atrial ﬁbrillation developed in collaboration with the Euro- pean Association for Cardio-Thoracic Surgery (EACTS): the task force for the diagnosis and management of atrial ﬁbrillation of the European Society of Cardiology (ESC) developed with the special contribution of the European Heart Rhythm Association (EHRA) of the ESC.  Eur   Heart  J.  2021;42(5):373-498.  Published correction appears in Eur  Heart  J. 2021;42(5):507. Published   correction   appears   in   Eur    Heart    J. 2021;42(5):546-547. Published correction appears in Eur Heart J. 2021;21;42(40):4194.

16

Saw et al

21. Reddy VY, Sievert H, Halperin J, et al. Percuta- neous left atrial appendage closure vs warfarin for atrial ﬁbrillation: a randomized clinical trial. JAMA. 2014;312(19):1988-1998.

22. Approval order for the Watchman  LAA closure technology (P130013). United States Department of Health & Human Services. Accessed July 21, 2021.

https://www.accessdata.fda.gov/cdrh_docs/pdf13/P13

0013A.pdf

23. Decision   memo   for   percutaneous   left   atrial appendage  (LAA)  closure  therapy  (CAG-00445N). Centers for Medicare & Medicaid Services. Accessed July    21,    2021.    https://www.cms.gov/medicare- coverage-database/details/nca-decision-memo.aspx? NCAId=281

24. Freeman JV, Varosy P, Price MJ, et al. The NCDR left atrial appendage occlusion registry. J  Am  Coll Cardiol. 2020;75(13):1503-1518.

25. Lip  GYH,  Nieuwlaat  R,  Pisters  R,  Lane  DA, Crijns  HJGM.  Reﬁning clinical  risk stratiﬁcation for predicting  stroke  and  thromboembolism  in  atrial ﬁbrillation using a novel risk factor-based approach: the  euro  heart  survey  on  atrial  ﬁbrillation.  Chest. 2010;137(2):263-272.

26. Boersma LV, Ince H, Kische S, et al. Efﬁcacy and safety   of   left   atrial   appendage   closure   with WATCHMAN in patients with or without contraindi- cation  to  oral  anticoagulation:  1-year  follow-up outcome   data   of  the   EWOLUTION   trial.   Heart Rhythm. 2017;14(9):1302-1308.

27. Landmesser U, Schmidt B, Nielsen-Kudsk JE, et al. Left atrial appendage occlusion with the AMPLATZER Amulet  device:  periprocedural  and  early  clinical/ echocardiographic  data  from  a  global  prospective observational   study.   EuroIntervention.  2017;13(7): 867-876.

28. Marroquin  L, Tirado-Conte G,  Praco/nR, et al.

Management and outcomes of patients with left atrial appendage thrombus prior to percutaneous closure. Heart. 2022;108(14):1098-1106.

29. Masoudi FA, Calkins H, Kavinsky CJ, et al. 2015 ACC/HRS/SCAI Left Atrial Appendage Occlusion De- vice Societal Overview: A professional societal over- view from the American College of Cardiology, Heart Rhythm Society, and Society for Cardiovascular Angi- ography and Interventions. Catheter Cardiovasc Interv. 2015;86(5):791-807.

30. Kavinsky CJ, Kusumoto FM, Bavry AA, et al. SCAI/ ACC/HRS institutional and operator requirements for left atrial appendage occlusion. Catheter  Cardiovasc Interv. 2016;87(3):351-362.

31. Zipes DP, Calkins H, Daubert JP, et al. 2015 ACC/

AHA/HRS  advanced  training  statement  on  clinical

cardiac electrophysiology (a revision of the ACC/AHA

2006 update of the clinical competence statement on

invasive electrophysiology studies, catheter ablation,

and cardioversion). Heart Rhythm. 2016;13(1):e3-e37.

32. Glikson M, Wolff R, Hindricks G, et al. EHRA/EAPCI

expert consensus statement on catheter-based left

atrial  appendage  occlusion   -  an   update.  Euro-

Intervention. 2020;15(13):1133-1180.

33. Wang DD, Geske JB, Choi AD, et al. Interventional

imaging for structural heart disease: challenges and

new frontiers of an emerging multi-disciplinary ﬁeld.

Struct Heart. 2019;3(3):187-200.

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

| 34. Wang DD, Forbes TJ, Lee JC, Eng MH. Echocar- diographic imaging for left atrial appendage occlusion: transesophageal echocardiography and intracardiac echocardiographic imaging. Interv Cardiol Clin. 2018;7(2):219-228. 35. Eng MH, Wang DD, Greenbaum AB, et al. Pro- spective, randomized comparison of 3-dimensional computed tomography guidance versus TEE data for left atrial appendage occlusion (PRO3DLAAO). Cath- eter Cardiovasc Interv. 2018;92(2):401-407. 36. Korsholm K, Berti S, Iriart X, et al. Expert recom- mendations on cardiac computed tomography for planning transcatheter left atrial appendage occlu- sion. J Am Coll Cardiol Intv. 2020;13(3):277-292. 37. Wang DD, Eng M, Kupsky D, et al. Application of 3- dimensional computed tomographic image guidance to WATCHMAN implantation and impact on early operator learning curve: single-center experience. J Am Coll Cardiol Intv. 2016;9(22):2329-2340. 38. Leipsic J, Nørgaard BL, Khalique O, et al. Core competencies in cardiac CT for imaging structural heart disease interventions: an expert consensus statement. J Am Coll Cardiol Img. 2019;12(12):2555- 2559. 39. Kaafarani M, Saw J, Daniels M, et al. Role of CT imaging in left atrial appendage occlusion for the WATCHMANTMdevice. Cardiovasc Diagn Ther. 2020;10(1):45-58. 40. Saw J, Fahmy P, Spencer R, et al. Comparing measurements of CT angiography, TEE, and ﬂuoros- copy of the left atrial appendage for percutaneous closure. J Cardiovasc Electrophysiol. 2016;27(4):414- 422. 41. Alkhouli M, Chaker Z, Alqahtani F, Raslan S, Raybuck B. Outcomes of routine intracardiac echo- cardiography to guide left atrial appendage occlusion. J Am Coll Cardiol EP. 2020;6(4):393-400. 42. Berti S, Pastormerlo LE, Santoro G, et al. Intra- cardiac versus transesophageal echocardiographic guidance for left atrial appendage occlusion: the LAAO Italian multicenter registry. J Am Coll Cardiol Intv. 2018;11(11):1086-1092. 43. Spencer RJ, DeJong P, Fahmy P, et al. Changes in left atrial appendage dimensions following volume loading during percutaneous left atrial appendage closure. J Am Coll Cardiol Intv. 2015;8(15):1935-1941. 44. Aminian A, Lalmand J, Tzikas A, Budts W, Benit E, Kefer J. Embolization of left atrial appendage closure devices: a systematic review of cases reported with the watchman device and the amplatzer cardiac plug. Catheter Cardiovasc Interv. 2015;86(1):128-135. 45. Gilhofer TS, Inohara T, Parsa A, et al. Safety and feasibility of same-day discharge after left atrial appendage closure. Can J Cardiol. 2020;36(6):945- 947. 46. Tan BE-X, Boppana LKT, Abdullah AS, et al. Safety and feasibility of same-day discharge after left atrial appendage closure with the WATCHMAN device. Circ Cardiovasc Interv. 2021;14(1):e009669. 47. Main ML, Fan D, Reddy VY, et al. Assessment of device-related thrombus and associated clinical out- comes with the WATCHMAN left atrial appendage closure device for embolic protection in patients with atrial ﬁbrillation (from the PROTECT-AF trial). Am J Cardiol. 2016;117(7):1127-1134. | 48. Sherwood M, Bliden KP, Ilkhanoff L, et al. Detailed thrombogenicity phenotyping and 1 year outcomes in patients undergoing WATCHMAN im- plantation: (TARGET-WATCHMAN) a case-control study. J Thromb Thrombolysis. 2020;50(3):484-498. 49. Vuddanda VLK, Turagam MK, Umale NA, et al. Incidence and causes of in-hospital outcomes and 30- day readmissions after percutaneous left atrial appendage closure: a US nationwide retrospective cohort study using claims data. Heart Rhythm. 2020;17(3):374-382. 50. Saw J, Nielsen-Kudsk JE, Bergmann M, et al. Antithrombotic therapy and device-related throm- bosis following endovascular left atrial appendage closure. J Am Coll Cardiol Intv. 2019;12(11):1067-1076. 51. Qamar SR, Jalal S, Nicolaou S, Tsang M, Gilhofer T, Saw J. Comparison of cardiac computed tomography angiography and transoesophageal echocardiography for device surveillance after left atrial appendage closure. EuroIntervention. 2019;15(8):663-670. 52. Korsholm K, Jensen JM, Nørgaard BL, Nielsen- Kudsk JE. Detection of device-related thrombosis following left atrial appendage occlusion: a compari- son between cardiac computed tomography and transesophageal echocardiography. Circ Cardiovasc Interv. 2019;12(9):e008112. 53. Al-Kassou B, Neikes F, Stock F, Buellesfeld L, Omran H. P6101 Incidence and persistence rates of iatrogenic atrial septal defects following left atrial appendage closure. Eur Heart J. 2017;38(suppl_1). ehx493.P6101. 54. Shah P, deFilippi CR, Mullins KE. Late-breaking science abstracts and featured science abstracts from the American Heart Association’s scientiﬁc sessions 2020 and late-breaking abstracts in resuscitation science from the Resuscitation Science Symposium 2020. Circulation. 2020;142(24):2293-2295. 55. Tzikas A, Gafoor S, Meerkin D, et al. Left atrial appendage occlusion with the AMPLATZER Amulet device: an expert consensus step-by-step approach. EuroIntervention. 2016;11(13):1512-1521. 56. Saw J, Tzikas A, Shakir S, et al. Incidence and clinical impact of device-associated thrombus and peri-device leak following left atrial appendage closure with the amplatzer cardiac plug. J Am Coll Cardiol Intv. 2017;10(4):391-399. 57. Pillarisetti J, Reddy YM, Gunda S, et al. Endocar- dial (Watchman) vs epicardial (Lariat) left atrial appendage exclusion devices: understanding the dif- ferences in the location and type of leaks and their clinical implications. Heart Rhythm. 2015;12(7):1501- 1507. 58. Nguyen A, Gallet R, Riant E, et al. Peridevice leak after left atrial appendage closure: incidence, risk factors, and clinical impact. Can J Cardiol. 2019;35(4): 405-412. 59. Viles-Gonzalez JF, Kar S, Douglas P, et al. The clinical impact of incomplete left atrial appendage closure with the Watchman Device in patients with atrial ﬁbrillation: a PROTECT AF (Percutaneous Closure of the Left Atrial Appendage Versus Warfarin Therapy for Prevention of Stroke in Patients With Atrial Fibrillation) substudy. J Am Coll Cardiol. 2012;59(10):923-929. 60. Alkhouli M, Du C, Killu A, et al. Clinical impact of residual leaks following left atrial appendage |
| --- | --- |

J A C C :   C A R D I O V A S C U L A R    I N T E R V E N T I O N S    V O L .    - ,   N O .   - ,   2 0 2 3

- ,  2 0 2 3 : - - -

| occlusion: insights from the NCDR LAAO registry. J Am Coll Cardiol EP. 2022;8(6):766-778. 61. Della Rocca DG, Horton RP, Di Biase L, et al. First experience of transcatheter leak occlusion with detachable coils following left atrial appendage closure. J Am Coll Cardiol Intv. 2020;13(3):306-319. 62. Alkhouli M, Alqahtani F, Kazienko B, Olgers K, SenGupta PP. Percutaneous closure of peridevice leak after left atrial appendage occlusion. J Am Coll Cardiol Intv. 2018;11(11):e83-e85. 63. Transcatheter leak closure with detachable coils following incomplete left atrial appendage closure procedures (TREASURE). ClinicalTrials.gov identiﬁer: NCT03503253. Accessed July 20, 2021. https:// clinicaltrials.gov/ct2/show/NCT03503253 64. Sleiman JR, Lewis AJ, Perez EJ, et al. Management of peri-device leak following left atrial appendage closure: a systematic review. Catheter Cardiovasc Interv. 2021;98(2):382-390. 65. Reddy VY, Doshi SK, Kar S, et al. 5-year outcomes after left atrial appendage closure: from the PREVAIL and PROTECT AF trials. J Am Coll Cardiol. 2017;70(24):2964-2975. 66. AMPLATZER Amulet left atrial appendage occluder trial (Amulet IDE). ClinicalTrials.gov identi- ﬁer: NCT02879448. Accessed July 20, 2021. https:// clinicaltrials.gov/ct2/show/NCT02879448 67. Williams T, Alsanjari O, Parker J, et al. Day-case percutaneous left atrial appendage occlusion-safety and efﬁcacy. Catheter Cardiovasc Interv. 2018;92(7): 1439-1443. 68. Kirchhof P, Benussi S, Kotecha D, et al. 2016 ESC guidelines for the management of atrial ﬁbrillation developed in collaboration with EACTS. Eur Heart J. 2016;37(38):2893-2962. | 69. Fauchier L, Cinaud A, Brigadeau F, et al. Device-related thrombosis after percutaneous left atrial appendage occlusion for atrial ﬁbrilla- tion. J Am Coll Cardiol. 2018;71(14):1528-1536. 70. Dukkipati SR, Kar S, Holmes DR, et al. Device- related thrombus after left atrial appendage closure: incidence, predictors, and outcomes. Circulation. 2018;138(9):874-885. 71. Kar S, Doshi SK, Sadhu A, et al. Primary outcome evaluation of a next-generation left atrial appendage closure device: results from the PINNACLE FLX trial. Circulation. 2021;143(18):1754-1762. 72. Coylewright M, et al. Comparative effectiveness of post-procedure medications following left atrial appendage occlusion: a DAPT analysis with WATCHMAN FLX. Presented at TVT. June 9, 2022. Chicago, USA. 73. Lakkireddy D, Thaler D, Ellis CR, et al. Amplatzer Amulet left atrial appendage occluder versus Watchman device for stroke prophylaxis (Amulet IDE): a randomized, controlled trial. Circulation. 2021;144(19):1543-1552. 74. Reddy VY, Möbius-Winkler S, Miller MA, et al. Left atrial appendage closure with the Watchman device in patients with a contraindication for oral anti- coagulation: the ASAP study (ASA Plavix feasibility study with Watchman left atrial appendage closure technology).J AmColl Cardiol. 2013;61(25):2551-2556. 75. Osman M, Busu T, Osman K, et al. Short-term antiplatelet versus anticoagulant therapy after left atrial appendage occlusion: a systematic review and meta-analysis. J Am Coll Cardiol EP. 2020;6(5):494- 506. 76. Shah PB, FGP Welt, Mahmud E, et al. Triage con- siderations for patients referred for structural heart disease intervention during the COVID-19 pandemic: |
| --- | --- |

Saw et al                 17

an ACC/SCAI position statement. J Am  Coll  Cardiol

Intv. 2020;13(12):1484-1488.

77. Phillips  KP,  Pokushalov  E,  Romanov  A,  et  al.

Combining Watchman left atrial appendage closure

and catheter ablation for atrial ﬁbrillation: multicentre

registry results of feasibility and safety during implant

and 30 days follow-up. Europace. 2018;20(6):949-

955.

78. Phillips  KP,  Romanov  A,  Artemenko  S,  et  al.

Combining left atrial appendage closure and catheter

ablation for atrial ﬁbrillation: 2-year outcomes from a

multinational registry. Europace. 2020;22(2):225-231.

79. aMAZE Study: LAA Ligation Adjunctive to PVI for

Persistent or Longstanding Persistent Atrial Fibrilla-

tion      (aMAZE).      ClinicalTrials.gov      identiﬁer:

NCT02513797.  Accessed  July  20,  2021.  https://

clinicaltrials.gov/ct2/show/NCT02513797

80. Lee RJ, Lakkireddy D, Mittal S, et al. Percutaneous

alternative to the Maze procedure for the treatment of

persistent or long-standing persistent atrial ﬁbrilla-

tion (aMAZE trial): rationale and design. Am Heart J.

2015;170(6):1184-1194.

81. WATCH-TAVR,  WATCHMAN  for   Patients  With Atrial  Fibrillation  Undergoing  Transcatheter  Aortic Valve    Replacement.   ClinicalTrials.gov    identiﬁer: NCT03173534.  Accessed  July  20,  2021.  https:// clinicaltrials.gov/ct2/show/NCT03173534

KEY WORDS  CHA2 DS2 -VASc, device-related      thrombus, embolic stroke prevention, left atrial appendage closure, non-valvular atrial

ﬁbrillation, peridevice leak

APPENDIX  For supplemental material, please see the online version of this paper.