# Guidelines

British Nuclear Medicine Society Clinical Guidelines for Gastric empty

Alp Notghia, Gregory Jamesb and Peter D. Hayc

Nuclear Medicine Communications 2023, 44:563–570

aDepartment of Physics and Nuclear Medicine, Consultant Physician in Nuclear  Medicine, Sandwell and West Birmingham NHS Trust, Birmingham,  bHead of    Nuclear Medicine Physics, Royal Stoke University Hospital, University Hospitals of North Midlands NHS Trust, Staffordshire and  c Medical Physics and Clinical  Engineering, Nottingham University Hospitals NHS Trust, Nottingham, UK

Correspondence to Alp Notghi, Department of Physics and Nuclear Medicine,

Consultant Physician in Nuclear Medicine, Sandwell and West Birmingham NHS Trust, City Hospital, Dudley Road, Birmingham B18 7QH, UK

E-mail: alp.notghi@nhs.net

Received 14 December 2022 Accepted 28 March 2023.

# Purpose

The purpose of this guideline is to assist specialists in nuclear medicine and radionuclide radiology in recom- mending,  performing,  interpreting  and  reporting  the results of gastric emptying studies. This guideline will assist individual departments to formulate their own local protocols. This does not aim to be prescriptive regard- ing technical aspects of individual camera acquisitions, which need to be developed in conjunction with the local Medical Physics Expert. These guidelines pertain only to adult patients.

# Background

Gastric  emptying  scintigraphy  is  a  safe,  non-invasive method for assessing the ability of the stomach to empty which has been used clinically for many years [1,2]. It is regarded as a ‘gold standard’ to assess gastric function and emptying of both solids and liquids allowing assessment of early, mid and late emptying, each of which may be altered by pathology.

The normal stomach serves as a temporary storage area for food. It physically and chemically breaks down food into smaller particles via muscular activity, hydrochlo- ric  acid  and  enzymes.  It  mixes  the  food  with  gastric secretions to form chyme, which it releases in a con- trolled manner through the pyloric  sphincter into the duodenum.

The  rate  of emptying  varies  with  a  variety  of fac- tors with liquids emptying more rapidly than solids [3]. Liquids tend to empty rapidly and solids empty slower  depending  on  several  factors. There  can  be some variation for normal gastric emptying even in the  same  subject.  Food  consistency,  volume,  calo- rific  content  and  chemical  composition  will  affect the rate of emptying, for example, a high-fat content slows  gastric  emptying.  Other  patient  factors  such as sex, smoking, time of day and time of the men- strual cycle in women may also affect gastric motility [4]. However, some of the evidence in these areas is contradictory.

Several drugs alter the rate of emptying [3] so a history of relevant medication should be taken prior to the test (which include prokinetic drugs such as erythromycin and metoclopramide and drugs slowing gastric empty- ing such as opiates and antispasmodics, see Appendix A). Where possible, drugs that affect the emptying rate should be stopped for at least four pharmaceutical half- lives prior to the test; however, this may not be possible or  advisable  and  should  be  tailored  to  the  individ- ual clinical context in consultation with the referring clinician.

The wide range of procedures used and the variability of gastric emptying make it essential that standard, vali- dated methods are applied [5]. For departments wishing to set up this study, the standard consensus meal used in the USA (Procedure guideline for adult solid meal gastric emptying  study—Society  of Nuclear  Medicine  guide- lines) is one option [6] as there is much experience with this method and it is relatively straightforward although the meal may not be suitable for all patients. This stand- ard consensus meal and its normal range are validated for the simplified method; where a standard egg-based meal is administered, and a single image at 3 or 4 h is acquired. Some departments consider this too simplistic and per- form a more detailed study. A recent survey of gastric emptying practice in the UK showed a wide variation in method with empirical normal ranges [7]. This guideline tries to address this, for a more consistent and palatable approach throughout the  UK. A more unified national protocol may help to develop a more consistent approach to gastric emptying studies with results applicable, trans- ferable and understandable across the centres, increasing the clinical confidence in the test. Where a department is wishing to use another method, then they should be encouraged to keep to that protocol where this is suita- ble for the local population but only if it has undergone a  process  of validation,  and  normal  ranges  are  locally established.

Gastric emptying studies require adequate clinical, medi- cal, physics and technical support to achieve good results.

0143-3636 Copyright © 2023 Wolters Kluwer Health, Inc. All rights reserved.                                                                        DOI: 10.1097/MNM.0000000000001699

564   Nuclear Medicine Communications     2023, Vol 44 No 7

Conditions that are investigated using gastric emptying scintigraphy include:

(1) Possible gastroparesis;

(2) Dumping syndrome;

(3) Possible gastroesophageal reflux;

(4) Endoscopy-negative functional dyspepsia;

(5) Persistent symptoms after gastric surgery;

(6) Follow-up scans to assess response to therapy;

(7) Motility studies for drug and physiological research.

# Contra-indications

Absolute

Allergy to the foodstuff (e.g. eggs, nuts, lactose, gluten) when these are included in a gastric emptying meal.

Relative

Pregnancy  and  breastfeeding.  If  low  activity  Tc-99m Colloid is used, there is no need to interrupt breastfeed- ing. However, as studies of gastric emptying are not for urgent life-threatening conditions, consideration should be given to delaying the investigation until after breast- feeding has ceased.

Diabetic patients require caution when fasting.

# Gastric emptying protocols

There are two validated protocols, which can be used for detailed gastric emptying studies:

A. Solid gastric emptying using standardised semi-solid meal followed by frequent imaging for at least 90 min. This method is most commonly used [7]. Here we describe a validated method with established normal ranges for an ‘egg & bread’ based meal, and also for a porridge-based meal as an alternative for those who are unable to have the ‘egg & bread’based meal (see section A).

A low-calorie liquid meal (typically 50–100 kcal in 100–200 ml) can be performed as an adjunct/ addi- tion to solid gastric emptying. It can be performed using Tc-99m, on a separate day to the solid gastric emptying, or using In-111 labelled meal if performed together with Tc-99m labelled solid meal as a dual isotope technique (see method A).

B. An alternative  stand-alone method that makes use of a high volume, and high-calorie liquid-based meal with established reference ranges [8,9]. This approach has been demonstrated to produce results equivalent to the use of a solid meal [10-12]. It also offers some operational advantages in terms of speed of prepa- ration,  meal  standardisation  and  being  of  shorter duration can potentially improve patient throughput (see method B). A recent publication describes prev- alent patterns of gastric emptying in a large cohort of patients using this method [13].

A gastric emptying study can be followed by a small bowel transit assessment where hourly images are continued up

to 6 h after ingestion. This can be followed by a separate colonic transit study.

Patient preparation is the same regardless of the method used.

Patient preparation

Fasting

Patients should be informed about what meal they will be expected to eat before the appointment so that they can inform the department of any allergies or religious/ cultural restrictions. Generally, the patient should fast at least 4 h prior to the procedure, ideally overnight. In dia- betic patients it is best to perform the test early in the morning, however, if the procedure is later in the day a light breakfast (e.g. dry toast) may be allowed to avoid hypoglycaemia. Diabetic patients should also bring their glucose  monitors  and  insulin  to  their  appointment  so that they can continue to self-monitor during the test. Smoking should also be avoided. All diabetic patients should have their blood glucose measured using a cali- brated glucose monitor. The gastric emptying test should only be carried out if the recorded blood glucose is less than 16 mmol per litre. With a high-volume high calorie liquid meal (method B), a further blood glucose measure- ment should be made 30 min into the study. If required, steps should be taken to control abnormal blood glucose levels.

Medication

Check food allergies and medications in advance. Where possible,  drugs  that  affect  the  emptying  rate  should be  stopped prior to the  study  (see Appendix A). The Administration   of   Radioactive    Substances   Advisory Committee (ARSAC) practitioner licence holder and the referring clinician should be consulted and authorisation obtained before stopping any of the medication. If it is not possible or advisable to stop some of the medication then this  should be considered when interpreting the study and clearly stated in the report.

Patient arrival

Check  patient  identity  and  fasting  status:  On  arrival, patient identity should be checked and the patient asked whether they have complied with the requirement to fast prior to the test, with the exception of diabetic patients as described above. All female patients in the age range of 12-55 years should be asked to confirm that they are not pregnant.

Image markers may be used to help align the images and correct for movement, this is particularly helpful if liquid meals are used. Radioactive fiducial markers may be pre- pared by using approximately 0.5 MBq of either Tc-99m or In-111, to match the radionuclide used in the test meal. These should be sealed and then attached to the skin so that they are visible in both anterior and posterior views.

British Nuclear Medicine Society Clinical Guidelines for gastric empty Notghi et al.    565

This can be done by fixing the markers to the skin to the right anterior and posterior costal margins of the patients prior to imaging, or by placing them on the left and right lateral aspect of the chest wall, to provide an anatomical reference during image analysis.

Effective radiation dose

The effective dose from 10 MBq of Tc-99m non-absorba- ble compounds in 0.3 mSv. The effective dose for 5 MBq of In-111  diethylene-triamine-pentaacetate  (DTPA)  is 1.6 mSv [14].

Technical imaging details:

| Camera Collimator Matrix size Image acqui- sition time Energy window | A dual-headed camera to allow simultaneous anterior and posterior acquisition. If a single-headed camera is used, then for static images consecutive anterior and posterior views need to be acquired. Tc-99m: Low energy general purpose collimator In-111: Medium energy general purpose collimator 128 × 128 or 256 × 256, whole field of view (un-zoomed acquisition) Static images are acquired for a minimum of 30 s for curve fitting, but longer acquisitions will give better count and image quality. For delayed images (3 or 4 h) longer acquisition times may be required. Tc-99m: 140 keV ± 10% In-111 : 171 keV ± 10% and 245 keV ± 10% |
| --- | --- |

# Method A: solid (±low-calorie liquid) gastric emptying study

The standard meal should be consumed within 5-10 min. A record of the time to start eating and the time to start scanning should be made.

Since  many  patients  complain  of sensations  such  as nausea, bloating and pain on eating, it is valuable to record patients’ symptoms during the study. This may be used to aid the interpretation of the investigation results.

For solid/ semi-solid gastric emptying studies, Tc-99m can be combined with any non-absorbable compound, which  does  not  dissociate  from  the  meal  (remaining in solid/semi-solid component of the meal) and is not absorbed  in  the  GI  tract. Colloids,  DTPA, and  MAA may be used depending on the meal. A low-calorie liq- uid gastric emptying study may be performed using a Tc-99m  labelled  non-absorbable  radiopharmaceutical, on a separate day. If a dual liquid/solid study is required, a simultaneous liquid study can be performed labelled with In-111 DTPA.

The  ARSAC  diagnostic  reference  level  for  Tc-99m non-absorbable compounds is up to 40 MBq. However, if intermittent static images are utilised lower activity would be appropriate (10-12 MBq). TheARSAC diagnostic ref- erence level for In-111-DTPA is 12 MBq, which may be used for dual-isotope liquid/solid studies. If a combined gastric emptying and colonic transit study are being car- ried out, then In-111 should be used (5-10 MBq).

The choice of the meal to which radiopharmaceuticals are labelled will depend on the protocol [5].

For  low-calorie  liquid  gastric   emptying  radiolabelled water or orange juice is often used. The radiopharmaceu- tical is mixed in 100-200 ml of liquid.

Two meals for solid/semi-solid gastric emptying studies, which are most commonly used in the UK [7], were stand- ardised for consistency across the centres and then nor- mal ranges were established for these meals [15]. These two alternative meals were chosen, as a single meal may not be suitable for all.

Scrambled Egg Meal

Scrambled egg meal is the recommended standard meal. Break two medium whole eggs into a bowl, add 99mTc- DTPA or colloid and mix with a fork. Cook in a micro- wave,  intermittently  mixing  to  achieve  the  required consistency and season with some salt. Serve with two slices of medium-cut bread; no butter or spread is used.

Porridge meal

This is the alternative meal for those who are unable to eat the scrambled egg meal for various reasons (allergies, vegetarian, gluten  sensitivity, etc.). Porridge with  soya milk has a similar composition to porridge made with whole-fat milk and may be substituted for patients with lactose intolerance or those who are vegan.

Mix  Tc-99m  labelled  DTPA  or  colloid  in  200 ml  of whole milk and add to 40 g of whole oats (gluten-free if required) and mix. Warm in microwave (following man- ufacturer’s instructions) and serve seasoned with some sweetener or salt if required.

See Appendix B for the composition and caloric content of the meals.

# Imaging procedure:

| Scanning mode | Solid and semi-solid meal studies: Simultaneous or sequential anterior/posterior planar images every 5–15 min. Less frequent imaging may affect the accuracy of some of the |
| --- | --- |
| Image time | measured parameters. Low-calorie liquid meal study: Dynamic images for the first 30 minutes (preferably with dual headed camera) followed by intermittent images as per solid meal protocol up to 3 hours. 2-min duration, static images for the first 2 h, then a 5-min |
| Patient position | image at 3 or 4 h. Preferably, erect position during the acquisition, with patient allowed to sit, in between the images. If images are acquired supine, then the patient must be allowed to stand or sit in between the images as gravity can have a significant effect on gastric emptying rate (normal range included in this guidance is not applicable if patient is lying down for the entire duration of the study). The stomach, distal oesophagus and small bowel should be included in the field of view. For the dynamic liquid study, patient should be standing in front of the camera heads. |
| Time of start | For solid and semi-solid meal, the images should start as soon as patient has consumed the whole meal. For low-calorie liquid stud- ies dynamic acquisition should start before the patent drinks the liquid, as in gastric dumping gastric emptying could be very rapid. |

# Considerations

If the patient is unable to consume the whole meal the results may be invalid, this should be noted down and highlighted for the attention of the reporter.

If the patient vomits, this should be noted down, and if possible the amount (radioactivity) calculated, either directly or indirectly (reduction in total camera counts). If vomiting occurs imaging may continue, but the study should be interpreted with caution. The act of vomiting may change the gastric emptying rate regardless of the amount vomited, and the study is often not reliable.

Given the potential for nausea and vomiting, it would be best practice to cover the detector placed in front of the patient, avoiding the air vents of the detector head.

This approach of acquiring several static images acquired up to 2 h provides many time points on the resulting time activity curve. These points provide the ability to fit curves, which can then yield functional parameters such as lag time, peak emptying rate, and time to peak emptying.

# Review and processing

Potentially there are up to 25 pairs of anterior and poste- rior static images for this study. For reliable results on all parameters, the maximum image interval is 15 min (min- imum of nine pairs of images in total). At each time point, the stomach counts are calculated using the geometric mean (GM) of anterior and posterior counts for the calcu- lation of gastric emptying parameters.

Region of interest (ROI) is drawn around the stomach, taking care to avoid small bowel and oesophageal activ- ity. The GM counts should be corrected for radioactive decay unless this has already been applied during the acquisition.

The  following  parameters  can  be  calculated: Stomach emptying  lag  time, half-emptying  time  (time  for  50% of stomach to empty), peak emptying time and rate and time to peak emptying, exponential half-life, and 3 or 4 h retention of the activity by the stomach.

Reporting

The study should be reported by the ARSAC license holder or suitably trained delegate with an understanding of both the clinical and technical issues involved.

Structured reports are recommended to include:

(1) Indication.

(2) Technical note: This should comment on any factor which may cause a source of error such as incomplete consumption  of  the  meal,  vomiting,  drugs,  move- ment, non-fasting, bowel activity overlying stomach, meal takes too long to eat or oesophageal problems.

(3) Description: This should include both positive and negative findings.

(4) Findings: Should  include correlation with previous studies or alternative modalities if appropriate. Any symptomatic data should be added to provide addi- tional information back to the referrer.

(5) Interpretation:

(a)   The clinical question should be answered in the report.

(b)   The pattern of emptying should be described (such as  different emptying rates at early and late phases). Significant variation should be com- mented on, including intra-gastric reflux.

(c)    Calculated parameters should be reported and compared to the normal ranges. Please see the normal range table for each meal.

Gastric emptying normal ranges (n = 42) for suggested meals [15]:

|  | Porridge |  | Scrambled egg sandwich |  |
| --- | --- | --- | --- | --- |
|  | Median | 95% CI | Median | 95% CI |
| Lag time (min) Half-emptying time (min) Peak emptying rate (%/min) Time-to-peak emptying (min) Exponential half-life (min) 3-h retention (%) | 3.2 43 1.4 17 31 1.3 | 0–13 18–73 4.5–0.8 0–38 21–57 <7 | 15 71 0.9 39 41 4.5 | 1–34 44–116 1.7–0.5 10–74 20–82 <17 |

Figure  1  shows  an  example  of  the  graphic  output for a patient with slow gastric emptying. The output includes  the  patient  gastric  emptying  exponential curve  plotted  against  the  normal  range with  calcu- lated  parameters.  The  graphic  output  is  a  useful visual assessment of the pattern of gastric emptying in relation to normal.

# Method B: High-volume, high-calorie liquid gastric emptying

Preparation of radiolabelled test meal Test meal

The   test   meal   consists   of  200 ml  Vanilla   Fortisip (Nutricia Ltd, Trowbridge, Wiltshire, UK) or equivalent, mixed with 200 ml water and either  10 MBq Tc-99m- DTPA (or 5 MBq In-111-DTPA for those patients who are referred for whole gut transit to include the addi- tional colonic transit imaging). The 400 ml radiolabelled drink is supplied to the patient in 5 disposable plastic cups, each containing  80 ml. See Appendix B  for  the food composition.

Meal intake

Controlled ingestion

Once in the imaging room, the patient should complete the first section of the sensation report form (appendix C) and then  drink one cup  (80 ml)  every  2 min  until the full 400 ml has been ingested (total time  10 min). If a patient is unable to ingest the complete drink, and/ or vomits, this should be recorded and reported to the clinician.

Fig. 1

Curve analysis with image points (green dots) and the fitted curve for a normal gastric emptying. The shaded red area is the normal range for gastric emptying.

Sensation

Recording patient sensation

Patients are asked to report their sensations of satiation, fullness, bloating, heartburn, nausea and epigastric pain on visual analogue scales (VAS) before the test drink is administered, directly after ingestion of the 400 ml liq- uid test meal, after 30 min and after 90 min of imaging (Appendix C).

Imaging procedure:

| Image acquisi- tion Patient position Time of start | Simultaneous or sequential anterior/posterior planar images every 5–15 min. The nominal times for image acquisition are 0, 5, 10, 15, 20, 30, 40, 50, 60, 75 and 90 min. after ingestion. By this time, if it appears that the stomach emptying is less than 50%, images should also be acquired at 105 and 120 min. after ingestion. The patient should be positioned standing erect, or seated, in front of the camera. Imaging should start immediately after the patient has ingested the last cup of radiolabelled Fortisip and completed the T0 sensation report (see above). |
| --- | --- |

Removal of fiducial markers

The  markers  should  be  removed  before  the  patient leaves the Nuclear Medicine clinic.

Image processing and data analysis Count data extraction

Image count data is extracted. GM images are created from the serial anterior and posterior views if acquired

simultaneously  on  a  dual-headed  camera. On  the  ini- tial GM image, a ROI is drawn around the stomach and abdomen/small bowel to ascertain the maximum counts (100% activity). Subsequently, ROIs are drawn around the stomach alone and a background ROI in all images. Counts within the stomach and small bowel ROIs are cor- rected for background counts and for radioactive decay. (Regional GM calculations can be used for sequential imaging where anterior and posterior views are acquired in series.)

Marker of early phase gastric emptying

The  corrected  count  rates  from  the  stomach  and  small bowel in the first image are used to calculate the 100% value for the study and to allow a calculation of the percentage of the drink remaining in the stomach in the initial image

Table 1     Reference ranges

|  | Parameter | Reference range in healthy volunteers |
| --- | --- | --- |
| Sensation Motor function | Fullness Score @ T0 Dyspepsia Score at any point (bloating, nausea, heartburn, epigastric pain) V0 T50 R[T50] | ≤60 ≤30 85–92% 38–56 min 0.75–1.10% per min |

T0 is time 0. V0 is percentage of drink remaining in the stomach in the initial image. T50 is the half-emptying time. R[T50] is rate of emptying at T50.

Fig. 2

In this example V0 is 87% (left Y-axis intercept), T50 is 65 min. (time point when percentage retention is half V0) and R[T50] is 0.70% per min (rate of emptying at T50 determined from the Stomach Emptying Rate plot against the right Y-axis). Using the reference ranges in Table 1, a report would indicate that early emptying is normal but that late emptying is slow.

(V0). To calculate this parameter, the percentage retention in the stomach should be plotted against time. An appropri- ate curve should be fitted to the data points. V0 is theY-axis intercept and can be used as a measure of the early phase gastric emptying. See Table 1 for reference ranges.

Half-emptying time

The gastric half-emptying time, T50, can be obtained from the plot by finding the time point when the percent retained in the stomach is half V0.

Marker of late-phase gastric emptying

The rate of gastric emptying at T50, R[T50], can be used as a measure of late-phase gastric emptying (the gradient of the emptying curve at T50). See Table  1 for reference ranges. Figure 2 shows an example time activity plot with results for V0, T50 and R[T50] shown below. In this example V0 is 87% (left Y-axis intercept), T50 is 65 min. (time point when percentage retention is half V0) and R[T50] is 0.70% per min (rate of emptying at T50 determined from the stomach emptying rate plot against the right Y-axis). Using the reference ranges in Table 1, a report would indicate that early emptying is normal but that late emptying is slow.

Reporting

The study may be reported by theARSAC license holder or  a  technical  report  may  be  provided  by  a  suitably trained operator. The report should include any technical notes, correlation with any previous studies, a statement of whether  patient  sensation was  normal  or  abnormal (Table 1), and a statement of whether motor function was slow, normal or fast in both the early (V0) and late phases of emptying R[T50]. Reference ranges were taken from an initial analysis of gastric sensory and motor function in healthy volunteers [8] (Table 1).

Auditable aspects

Adherence to departmental scan acquisition, processing and  reporting  may  be  made  against  these  guidelines. Gastric emptying can be controversial and further work and audit is needed in this area.

The protocol for the method used should be reviewed every 3 years to ensure that it is still valid and that there has been no drift from established practice.

Acknowledgements

This  guideline  must  be  read  in  conjunction with  the BNMS Generic Guidelines.

Conflicts of interest

There are no conflicts of interest.

References

1   Griffiths GH, Owen GM, Kirkman S, Shields R. Measurement of the rate of gastric emptying using Chromium-51. Lancet 1966; 1:1244–1245.

2   Malmud LS, Fisher RS, Knight LC, Rock E. Scintigraphic evaluation of gastric emptying. Semin Nucl Med 1982; 12:116–125.

3   Feldman M, Lee EL. Pathophysiology/diagnosis/management. Sleisenger and Fordtran’s gastrointestinal and liver disease. s.l.: Saunders Elsevier;

2010. Vol. 1. ISBN: 978-1-4160-6189-2.

4   Brennan IM, Feltrin KL, Nair NS, Hausken T, Little TJ, Brennan IM, et al.

Effects of the phases of the menstrual cycle on gastric emptying, glycemia, plasma GLP-1 and insulin, and energy intake in healthy lean women. Am J  Physiol Gastrointest Liver Physiol 2009; 297:G602–G610.

5   Frier M, Perkins A. Radiopharmaceuticals and the gastrointestinal tract. Eur J Nucl Med 1994; 21:1234–1242.

6   Donohoe KJ, Maurer AH, Zeissman HA, Urbain JC, Royal HD, Martin-Comin

J. Procedure guideline for adult solid-meal gastric-emptying study 3.0. s.l.:

Soc Nucl Med 2009; 37:196–200.

7   Hansrod S, Notghi A. National survey of gastric emptying studies in the United Kingdom. Nucl Med Commun 2018; 39:957–959.

8   Parker HL, Tucker E, Hoad CL, Pal A, Costigan C, Hudders N, et al.

Clinical assessment of gastric emptying and sensory function utilizing   gamma scintigraphy: establishment of reference intervals for the liquid  and solid components of the Nottingham test meal in healthy subjects.

Neurogastroenterol Motil 2017; 29:e13122.

9   Parker HL, Tucker E, Hoad CL, Pal A, Costigan C, Hudders N, et al.

Development and validation of a large, modular test meal with liquid and

solid components for assessment of gastric motor and sensory function by non-invasive imaging. Neurogastroenterol Motil 2016; 28:554–568.

10   Ziessman HA, Chander A, Clarke JO, Ramos A, Wahl RL. The added

diagnostic value of liquid gastric emptying compared with solid emptying alone. J Nucl Med 2009; 50:5.

11   Solnes L, Sheikhbahaei S, Zeissman HA. EnsurePlus as an alternative to the   standardized egg gastric-emptying meal. Clin Nucl Med 2019; 44:459–461.

12   Ziessman H, Okolo P, Mullin G, Chander A. Liquid gastric emptying is

often abnormal when solid emptying is normal. J Clin Gastroenterol 2009;

43:639–643.

13    Hay PD, Corsetti M, Tucker E, Fox MR, Perkins A. Tolerability, gastric emptying patterns, and symptoms during the Nottingham Test Meal in 330 secondary

care non-diabetic dyspeptic patients. Neurogastroenterol Motil 2022:e14454.

14   Administration of Radioactive Substances Advisory Committee

(ARSAC). ARSAC notes for guid]ance: good clinical practice in nuclear medicine. [Online] 2023. https://www.gov.uk/government/publications/  arsac-notes-for-guidance.

15   Hansrod S, James G, Notghi A, Croasdale J, O’Brien J, Thomson W.

Gastric emptying: methodology and normal ranges for two commonly used meals in the UK. Nucl Med Commun 2020; 41:636–650.

Appendix A

Drugs affect gastric motility

There is a number of medication, which may affect gas- tric emptying by either stimulating or slowing the gastric emptying. The  decision to  stop medication for gastric emptying studies should be made in conjunction with the referring clinician, as in some cases it may not be advisable to stop these even temporarily. The following list is a comprehensive, but not exhaustive list of these drugs. Most of these drugs have relatively short half-lives. If possible and safe, they could be stopped at least 2 days prior to the gastric emptying study.

Drugs that stimulate gastric motility:

(1)   Metoclopramide: (Reglan, Metozolv ODT, Gimoti) It  is  a  dopamine  antagonist.  Indirect  action  on smooth   muscles   through   increased   acetylcho- line secretion. It has a half-life is 4 h. It does not increase gastric acid secretion.

(2)   Bethanechol:  Directly  stimulates  the  muscarine receptors of gastrointestinal smooth muscles, and also increases gastric acid secretion

(3)   Domperidone:   Dopamine   antagonist,   probably works by blocking alpha 1-adrenoceptors, does not increase gastric acid secretion.

(4)   Clebopride: Dopamine antagonists and stimulator of the intramural cholinergic system.

(5)   Trimebutine: Spasmolytic.

(6)   Cisapride (Propulsid): serotonin-4 (5-HT4) recep- tor agonists that stimulate GI motility. It acts by facilitating  the  release  of acetylcholine  from  the myenteric plexus. Its half-life is 10 h.

(7)   Tegaserod (Zelnorm): serotonin-4 (5-HT4) recep- tor agonists that stimulate GI motility.

(8)   Lubiprostone (Amitiza, Lubilax, Lubowel).

(9)   Prucalopride (Resolor).

(10)   Erythromycin: Motilin agonist. Accelerates gastric emptying.

Drugs that decrease gastric motility:

(1) Opioids:Includes morphine,meperidine,methadone):

(a)   Oxycodone (OxyContin, Roxicodone, and Percocet);

(b)   Hydrocodone (Vicodin, Norco, Lortab);

(c)   Morphine (MS Contin and Kadian);

(d)   Hydromorphone (Dilaudid and Exalgo);

(e)   Fentanyl   (Actiq,  Fentora,  Duragesic,  Subsys, Abstral, and Lazanda);

(f)    Meperidine/pethidine (Demerol);

(g)   Methadone (Metharose, Physeptone).

(2) Opioid analogues with more motility effect:

(a)   Diphenoxylate  (in  combination  with  atropine marketed   as   Lomotil):  Diphenoxylate   is   an opium derivative, and rapidly metabolised (by ester hydrolysis) to biologically active metabolite ‘difenoxylic acid.’

(b)   Loperamide hydrochloride (Imodium): Depresses longitudinal and circular muscle activity.

(3) 5-HT receptor agonists:

(a)   Alosetron: (Lotronex) is a 5-HT3 receptor antag- onist, reducing GI motility.

(b)   Ondansetron  (Ondemet, Setofilm, Zofran)  is  a 5-HT3 receptor antagonist, reducing GI motility.

(c)   Sumatriptan  (Imigran,  Migraitan)  is  a  5-HT1 receptor antagonist, reducing GI motility. Used to treat migraine headache.

(d)   Sildenafil  (Viagra,  Revatio)  slows  gastric  emp- tying when used in high dose (such as used for relaxation of oesophageal muscles in oesopha- geal motility disorders).

Other medication that may affect gastric emptying:

(1) Calcium channel blockers:

(a)   Amlodipine (Norvasc),

(b)   Diltiazem (Cardizem, Tiazac),

(c)   Nifedipine (Adalat CC, Procardia), (d)   Verapamil (Calan, Verelan),

(e)   Felodipine,

(f)    Isradipine,    (g)   Nicardipine.

(2) Tricyclic antidepressants:

(a)   Amitryptiline (Elavil),

(b)   Nortriptyline (Aventyl, Pamelor), (3) Dopamine agonists:

(a)   Bromocriptine (Parlodel), (b)   Cabergoline (Dostinex),   (c)   Apomorphine (Apokyn),   (d)   Pramipexole (Mirapex),    (e)   Ropinirole (Requip),

(f)    Pergolide (withdrawn from the USA market).

(4) Alpha-2-adrenergic agonists:

(a)   Clonidine (Catapres),    (b)   Methyldopa (Aldomet),

(c)   Tizanidine (Zanaflex) - used as a muscle relaxer, (d)   Clonidine (Kapvay) - used to treat ADHD,

(e)   Guanfacine (Intuniv) - used to treat ADHD,

(f)    Lofexidine  (Lucemyra)  - used  to  treat opiate withdrawal.

(5) Muscarinic cholinergic receptor antagonists:

(a)   Atropine (Atropine),

(b)   Belladonna alkaloids,

(c)   Benztropine mesylate (Cogentin), (d)   Clidinium (in Librax),

(e)   Cyclopentolate (Cyclogyl), (f)    Darifenacin (Enablex),

(g)   Dicyclomine (Bentyl),  (h)   Fesoterodine (Toviaz).

(6) Phenothiazines (Antipsychotic):

(a)   prochlorperazine (Compazine,Compro,Procomp), (b)   chlorpromazine (Promapar, Thorazine),

(c)   fluphenazine (Permitil, Prolixin), (d)   perphenazine,

(e)   trifluoperazine (Stelazine), (f)    thioridazine (Mellaril).

(7) Glucagon-like peptide (GLP)-1 agnosis and amylin analogues (for treatment of diabetes):

(a)   Albiglutide (Tanzeum), (b)   Dulaglutide (Trulicity), (c)   Exenatide (Byetta),

(d)   Extended-release exenatide (Bydureon), (e)   Liraglutide (Victoza),

(f)    Lixisenatide (Adlyxin),

(g)    Semaglutide (Ozempic, Rybelsus).

(8)   Cyclosporine (Immunosuppressant agent).

(9)   Progesterone.

(10)   Octreotide.

(11)   Benzodiazepine.

(12)   Phentolamine.

Appendix B

Composition of various meals used in gastric emptying:

|  | Scrambled egg and bread | 40 g oatmeal + 200 ml full fat milk | 40 g oatmeal + 200 ml soya milk | 200 ml vanilla fortisip + 200 ml water |
| --- | --- | --- | --- | --- |
| Calories | 270 | 280 | 230 | 300 |
| Carbohy- drates | 24 | 34 | 28 | 37 |
| Protein | 18 | 11 | 11 | 12 |
| Fat | 10 | 10 | 7 | 12 |

Appendix C

Sensation report form: